## HbA1c only
S.no | Product Name | [17 empty] | Technical | [5 empty] | Procedure | [6 empty] | Performance |  |  |  | Analyzer | [5 empty] | Resources | Cost (USD) | Availability | [10 empty] | Supplier | [10 empty] | Cost
|  | Manufacturer | SRA Approvals |  |  |  | Certifications |  | Method |  | Test Design |  |  | Device Design |  | Dimensions in mm | Device weight in kg | Additional Test | Use | User | Type of Sample | Volume | Duration | Reporting Units | Sample Prep | Handling | Equipment | Not provided | Control | Calibration | Storage temperature and shel life | Reference Range | Limitation | Precision | Interference | Data Export | Operating temperature | Humidity | Power | Display | Memory |  |  | Bangladesh | Bhutan | Republic of Korea | India | Indonesia | Maldives | Myanmar | Nepal | Sri Lanka | Thailand | Timor-Leste | Bangladesh | Bhutan | Republic of Korea | India | Indonesia | Maldives | Myanmar | Nepal | Srilanka | Thailand | Timor-Leste | Bangladesh | Bhutan | Republic of Korea | India | Indonesia | Maldives | Myanmar | Nepal | Sri Lanka | Thailand | Timor-Leste
1 | CERA STAT 2000 | Green Cross MEDIS Corp | CE-IVD |  |  |  | IFCC | NGSP | Boronate affinity |  | Cartridge |  |  | Benchtop |  | 178 x 195 x 75 | 0.73 | No | CERA-STAT 2000 analyzer is designed to measure the CERASTAT HbA1c Test Kit . The colored responses of the tests are measured by spectral reflectance in three parts of the visible spectrum. For In Vitro Diagnostic Use. | Professional use | Capillary whole blood, anticoagulated venous whole blood | 5 µL | <3 min | IFCC(mmol/mol)/NGSP/DCCT (%) values | Equilibrate reagents to room temperature. Collect the sample using the capillary tube. Put the capillary tube into the R1/ Reagent tube and mix throughly. Stand the tube for 2 minutes for the incubation. Apply 25 μL of the reaction mixture to the membrane of the cartridge. After 10 seconds, apply 25 μL of the R2/Reagent to the membrane of the cartridge. | Open the tray and put the cartridge in the tray | Analyzer: Cera Stat 2000, battery, power adaptor, printer paper roll, calibration plate, user manual. Test: tubes of Reagent 1, cartridges and bottle of Reagent | Standard blood collection equipment, pipette, pipette tips | CERA STAT 2000 Control kit contains two ready-to-use liquid controls | Auto calibration | At 2–8°C, 12 months | 3.0–15.0% / 9–140 mmol/mol | NA | <3.1% | An Hb-concentration lower than 10g/dL or higher than 20g/ Dl can cause inaccurate test results. The test results are not affected by albumin, ascorbic acid, biliorubin, flucose, lipid. | USB, RS-232 interface | 15–35°C | 15–75% | AC mains power no battery | 3.5 inch colour touch screen | 300 tests | CERA-STAT 2000 HbA1c ¦ Green Cross Medical Science Corp. (omnia-health.com) | Promoting Management and Leadership - HealthManagement.org ¦ Glycated hemoglobin analyzer CERA-STAT™ 2000 CERAGEM Medisys | NA | NA | Available | Available | NA | NA | NA | NA | NA | Available | NA | NA | NA | Green Cross Medical Science Corp. ¦ Medical Manufacturer Directory (omnia-health.com) Ceragem Medisys Inc. 3-2 Jeongchon-ri, Seonggeo-eup, Seobuk-gu, Cheonan, Chungcheongnam, 331833, Korea South | CERAGEM CERA-STAT™ 2000 for Sale - Bimedis - India | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
2 | clover A1c plus | OSANG | CE-IVD | FDA |  |  | IFCC | NGSP | Boronate affinity |  | Cartridge |  |  | Benchtop |  | 200 x 200 x 139 | 1.4 | No | The CLOVER A1c Plus is an IVD device for measuring Hemoglobin A1c in whole blood. The system is designed to help in controlling diabetes | Professionals in laboratories, clinics and hospitals | Capillary whole blood, anticoagulated venous whole blood | 4 µL | 5 min | IFCC(mmol/mol)/NGSP/DCCT (%) values | If refrigerated, equilibrate to room temperature before use. The sample collection device is in the cartridge. Remove it. Collect blood from finger. | Insert the test cartridge into the analyzer. Press the sample collector into the cartridge. | Analyzer: Clover A1c Plus, instruction manual, DC 9 volt, 2.0A adapter, fan filters, daily check cartridge, monthly check cartridge. Tests: cartridge with sample collector | Lancet, wipes | No control solutions provided. Two level control solutions from other manufacturers can be used. | NA | 2–32°C, 18 months | 4.0–14.0% / 20–130 mmol/mo | The CLOVER A1c Plus assay gives accurate and precise results over a range of total Haemoglobin of 7 to 20 g/dL. Patients with severe anaemia may have Haemoglobin concentrations lower than 7 g/dL, and patients with polycythaemia may have Haemoglobin concentrations above 20 g/dL | <3% | NA | Communication with PC by cable (USB to Serial port) | 10–40°C | 10–90% | DC 9V, 2A Adapter No battery | LCD | 1000 tests | HbA1C Analyzer (Clover A1c) ¦ tradekorea | 596 (INR 50000) | NA | NA | Available | Available | Available | NA | NA | NA | NA | NA | NA | NA | NA | OSANG Healthcare (osanghc.com) | 1. Blood Glucose Meter Supplier,Digital Blood Glucose Meter Exporter,India (kannuimpex.com) | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | INR 50000 | NA | NA | NA | NA | NA | NA | NA
3 | clover A1c | OSANG | CE-IVD | FDA |  |  | IFCC | NGSP | Boronate affinity |  | Cartridge |  |  | Benchtop |  | 200 x 198 x 139 | 1.4 | No | Clover A1c is a type of test used to measure hemoglobin A1c (HbA1c) levels in the blood. This test provides insight into a person's average blood sugar levels over the past two to three months, which is crucial for managing diabetes. | Professional use | Capillary whole blood, anticoagulated venous whole blood | 4 µL | 5 min | IFCC(mmol/mol)/NGSP/DCCT (%) values | If refrigerated, equilibrate to room temperature before use. The sample collection device is in the cartridge. Remove it. Collect blood from finger. | Open the analyzer, Insert the cartridge, insert the reagent pack into the cartridge in the analyzer, close the lid of the analyzer. | Analyzer: Clover A1c, Adapter, power cable, user manual; Test kit: Reagent pack, cartridge, sample collector. | Lancet, wipes | NA | NA | 2–30°C | 4.0–14.0% / 20–130 mmol/mo | NA | NA | NA | NA | 17–32°C | NA | DC 9V 2.0A | NA | 200 tests | Glycated hemoglobin – OSANG Healthcare (osanghc.com) | NA | NA | NA | NA | Available INR 75000 | Available | NA | NA | NA | NA | NA | NA | NA | NA | NA | Blood Glucose Meter Supplier,Digital Blood Glucose Meter Exporter,India (kannuimpex.com) | Sell Clover A1c Self A1c Meter Cheap Hb Check Tool (medicalogy.com) | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | IDR 20556250 https://www.medicalogy.com/p/beli-a1c-meter/clover-a1c-self-alat-cek-hba1c-AAG00001 | NA | NA | NA | NA | NA | NA
4 | The QDx A1c | DiaSys India | CE-IVD | FDA |  |  | IFCC | NGSP | Boronate affinity |  |  |  | Cassette | Benchtop |  | 205 x 220 x 150 | 1.6 | No | The QDx A1c Advance System is primarily used for the quantitative determination of Hemoglobin A1c (HbA1c) and estimated Average Glucose (eAG) levels in individuals | Professionals in laboratories, clinics and hospitals | Capillary whole blood, anticoagulated venous whole blood | 3.5 µL | 4 min 30sec | IFCC(mmol/mol)/NGSP/DCCT (%) values | NA | Insert the reagent pack into the cassette in the analyzer. The test starts automatically once the lid is closed | Analyzer, cassette, reagent pack | NA | NA | Auto calibrated test cassettes | 17–32°C | 4.0-15.0%/ 9-140mmol/mol | NA | <3% | NA | USB | 17-75% | 15-75% | DC 12V-3.33 A | 3.5 inch colour touch screen | 5000 tests | QDx A1c Advance - DiaSys India | 894 (INR 75000) | NA | NA | NA | Available INR 75000 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | INR 75000 | NA | NA | NA | NA | NA | NA | NA
5 | Pocketchem A1C | Arkray |  |  |  | ISO 13485, Japan | IFCC | NGSP | Boronate affinity |  | Cartridge |  |  | Benchtop |  | 205 x 135 x 95 | 0.7 | No | PocketChem A1c is a fast, accurate and simple glycohemoglobin analyzer with ergonomic design. | Professional use | Capillary whole blood, anticoagulated venous whole blood | 4 µL | 4 min | IFCC(mmol/mol)/NGSP/DCCT (%) values | Collect sample and insert sample stick halfway into cartridge | Insert into cartridge and close the lid | Analyzer,Catridges,reader,sample collector | NA | NA | Use of barcode reader to scan calibration data | NA | 4.0-15.0%/ 9-140mmol/mol | NA | <3% | NA | USB | 18–30°C | 20-80% | AC100V-240V +/- 10% 50/60Hz | Touch screen | 7000 tests | PocketChem A1c ¦ ARKRAY, Inc. | 1192 (INR 100000) | NA | NA | NA | Available INR 100000 , INR 94999 , INR 91332 | Available | NA | NA | NA | NA | NA | NA | NA | NA | NA | 1. Sri Ram Diagnostics, 2. https://www.hospitalstore.com/compact-glycohemoglobin-analyzer-pocketchem-a1c/ 3. https://www.colmed.in/arkray-compact-glycohemoglobin-analyzer-pocketchem-a1c.html.html | ARKRAY, Inc. | NA | NA | NA | NA | NA | NA | NA | NA | NA | INR 75000 | NA | NA | NA | NA | NA | NA | NA
6 | HemoCue HbA1c 501 | HemoCue | CE-IVD |  |  |  | IFCC | NGSP | Boronate affinity |  | Cartridge |  |  | Benchtop |  | 198 x 217 x 136 | 1.3 | No | The HemoCue® HbA1c 501 System provides a convenient method for measuring the percentage of hemoglobin A1c (HbA1c %) in both capillary and anticoagulated venous whole blood samples. The test is for point of care use to monitor glycemic control in patients with diabetes mellitus. | Professionals in laboratories, clinics and hospitals | Capillary whole blood, anticoagulated venous whole blood | 4 µL | 5 min | IFCC(mmol/mol)/NGSP/DCCT (%) values | If refrigerated, equilibrate to room temperature before use. The reagent pack is in the cartridge. Remove it and shake gently. Apply the reagent pack to the fingertip to collect blood from finger. | Insert the test cartridge into the analyzer. Press the reagent pack into the cartridge | Analyzer: Hemocue® HbA1c 501 analyzer, operating manual, exclusive power adapter, fan filters, daily check cartridge, monthly check cartridge. Tests; cartridge with reagent pack | NA | Built-in ”self test”. Check Cartridge. System can be verified using liquid controls | Factory calibrated. Lot-specific calibration | At 2–32°C, 18 months | 4.0–14.0% / 20–130 mmol/mo | The HemoCue HbA1c 501 assay gives accurate and precise haemoglobin results in the range 7 to 20 g/dL. Patients with severe anaemia may have haemoglobin concentrations below 7 g/dL, and patients with polycythaemia may have haemoglobin concentrations above 20 g/dL. Patients known to have these conditions should be tested with another method for determination of HbA1c % | < 3% | Studies confirm that no interference for the variants C, D, E, F and S were obtained, but HbF values above 20% may cause unreliable results | USB serial port to PC | 17–32°C | 10–90% | 9 V DC / 2 A | yes | 1000 tests | https://www.hemocue.com/en/ | 914.26 (INR 76700) | NA | NA | NA | Available INR 76700 | NA | NA | NA | NA | NA | Available | NA | NA | NA | NA | 1.https://www.hemocue.in/ 2.https://www.saaskin.in/ | NA | NA | NA | NA | NA | https://www.hemocue.in/en-in/health-areas/diabetes/launch-of-hemocue-hba1c/thailand-distributor | NA | NA | NA | NA | INR 1,00,000 , INR 94999 , INR 91332 | NA | NA | NA | NA | NA | NA | NA
7 | HumaMeter | HUMAN Diagnostics | CE-IVD |  |  |  | IFCC | NGSP | Boronate affinity |  | Cartridge |  |  | Benchtop |  | 205 x 95 x 135 | 0.7 | No | Quantitative measurement of glycated haemoglobin (HbA1c) and ratio (% or mmol/mol) of HbA1c to total haemoglobin (Hb). The HumaMeterA1c system is intended for the diagnosis of diabetes mellitus and for monitoring longterm glycaemic control in patients diagnosed with diabetes | Professional use | Capillary whole blood, anticoagulated venous whole blood | 4 µL | 4 min | IFCC(mmol/mol)/NGSP/DCCT (%) values | Equilibrate to room temperature. Scan lot calibration. Remove cartridge foil. IInsert cartridge into analyzer. Press reagent bead into cartridge. Insert sample stick into cartridge. | Close analyzer lid | Analyzer: HumaMeter, user manual, barcode scanner, power supply. Tests: test cartridges, sample sticks, instructions for use, calibration barcode | NA | HumaMeter A1c Control Kit contains level 1 & level 2 liquids | Factory calibrated. Lot-specific calibration | At 2–8°C, 12 months At RT, 90 days | 4.0–15.0% / 20–140 mmol/mol | NA | <3% at 7 %A1c | No interference from Hb variants other than those affecting erythrocyte lifespan | USB port to PC | 18–30°C | <85% non-condensing | 110...240 v AC, 24 V DC / 1.25A | LCD Screen 70 x 40 mm | 7000 tests | https://www.human.de/about-human/human-worldwide | 1492.5 (125000) | NA | NA | NA | Available INR 1,25,000 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | https://www.cadenamedical.com/ https://www.indiamart.com/proddetail/hba1c-analyzer-9079685062.html | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | INR 76700 | NA | NA | NA | NA | NA | NA | NA
8 | Quo-Lab | EKF Diagnostics | CE-IVD |  |  |  | IFCC | NGSP | Boronate affinity |  | Cartridge |  |  | Benchtop |  | 205 x 135 x 95 | 0.7 | No | The Quo-Lab A1C System is intended for the in vitro quantitative determination of glycated hemoglobin (HbA1c) in whole blood. The Quo-Lab A1C System is indicated in the management and treatment of diabetes mellitus and for monitoring long term glycemic control in patients diagnosed with diabetes | Clinics and laboratories in settings that demand low cost of operation and ease of use | Capillary whole blood, anticoagulated venous whole blood | 4 µL | 4 min | IFCC(mmol/mol)/NGSP/DCCT (%) values | Insert cartridge into analyzer. Press reagent bead into cartridge. Insert sample tool into cartridge and snap off handle. | Collect blood sample. Insert sample collector into the cartridge. | Analyzer: Quo-Lab Analyzer, barcode scanner, mains power cable, power supply, user manual. Test: test cartridges, blood collectors, calibration barcode, instructions for use | Fingertip lancets | Quo-Lab A1c Control Kit contains two level control solutions | Factory-calibrated Lot-specific calibration | At 2–8°C, 12 months At 15–25°C, 30 days | 4.0–15% / 20–140 mmol/mol | For haemoglobin concentrations outside the range of 6.5 g/dL to 20.4 g/dL (4.0 mmol/l to 12.7 mmol/l) or for HbA1c values outside the range of 4.0 % to 15.0 % DCCT (20 mmol/mol to 140 mmol/mol IFCC), a result will not be reported and an error message will be displayed. | <3% @7%A1c | Unaffected by Hb variants which do not result in reduced erythrocyte life span | via USB to PC | 18–30°C | 10–80% | Separate AC to DC mains adapter Input: 100–240 V AC, 50-60 Hz, 30 W Output: 24 V DC, 1.25A | Integrated screen | 7000 tests | https://www.ekfdiagnostics.com/point-of-care/diabetes/quo-lab/ | 596 (INR 50000) | NA | NA | NA | Available | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 1.Rapid Card and Biochemistry Analyzer Reagents Manufacturer ¦ Raj Biosis Private Limited, Jaipur 2.https://www.ozonebio.com/india 3. https://www.hospitalstore.com/quo-lab-hba1c-analyzer/ | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | INR 50000, INR 100000, INR 97000 | NA | NA | NA | NA | NA | NA | NA
9 | Quo-Test | EKF Diagnostics | CE-IVD | FDA |  |  | IFCC | NGSP | Boronate affinity |  | Cartridge |  |  | Benchtop |  | 205 x 135 x 205 | 1.3 | No | The Quo-Test A1C System is intended for the in vitro quantitative determination of glycated hemoglobin (HbA1c) in whole blood. The Quo-Test A1C System is indicated in the management and treatment of diabetes mellitus and for monitoring long term glycemic control in patients diagnosed with diabetes | For professional use only | Capillary whole blood, anticoagulated venous whole blood | 4 µL | 4 min | IFCC(mmol/mol)/NGSP/DCCT (%) values | Collect blood sample. Insert sample collector into the cartridge. | Insert cartridge into analyzer | Analyzer: Quo-Test Analyzer, barcode scanner, mains power cable, power supply, user manual. Test: test cartridges, blood collectors, calibration barcode, instructions for use | Fingertip lancets | Quo-Test A1c Control Kit contains two level control solutions | Factory-calibrated Lot-specific calibration | At 2–8°C, 12 months At 15–25°C, 30 days | 4.0–15% / 20–140 mmol/mol | For haemoglobin concentrations outside the range of 6.5 g/dL to 20.4 g/dL (4.0 mmol/l to 12.7 mmol/l) or for HbA1c values outside the range of 4.0 % to 15.0 % DCCT (20 mmol/mol to 140 mmol/mol IFCC), a result will not be reported and an error message will be displayed. Consult the Quo-Test Analyzer System User Manual for further details. | <3% @7%A1c | Unaffected by labile glycated haemoglobin and the following haemoglobin variants: Hb AS, Hb AC, Hb AD, Hb AE, Hb AJ, Hb SS, Hb CC, Hb SC, and elevated foetal haemoglobin (up to 6 %). | via USB to PC | 18–30°C | 10–80% | Separate AC to DC mains adapter Input: 100–240 V AC, 50-60 Hz, 30 W Output: 24 V DC, 1.25A | Blue LCD backlit monochrome display 128x64 pixels. Visible area 70x39mm | 7000 tests | https://www.ekfdiagnostics.com/point-of-care/diabetes/quo-test/ | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
10 | A1CNOW®+ | PTS diagnostics | CE-IVD | FDA | CLIA waived |  | IFCC | NGSP |  | Immunoassay | Cartridge |  |  |  | Handheld | 177.8 x 101.6 x 228.6 | 0.49 | No | The A1CNow®+ system provides healthcare professionals with a fast and easy way of obtaining accurate A1C results with just a fingerstick. | For professional use only | Capillary whole blood, anticoagulated venous whole blood | 5 µL | 5 min | IFCC(mmol/mol)/NGSP/DCCT (%) values | Equilibrate to room temperature. Apply collection device to a shaker, shake briefly and apply to cartridge | Blood collector measuring 5 uL blood sample (located in pouch 1). After sample application is applied, the test cartridge will activate the 5 minute countdown | (1) A1c Now+ Monitor, (10) Single-Use Cartridges, (10) Sample Dilution Kits, (1) Product Insert / Procedure Guide | Fingertip lancets | Over 50 internal chemical and electronic quality control checks, including potential hardware and software errors, and potential test strip errors. The analyzer has been programmed to report an error if these quality checks are not passed. Liquid control solutions and reagent test strip available. | Analyzer is factory-calibrated. Lot-specific calibration for tests | At 2–8°C, until expiration date At 18–28°C, 4 months | 4.0–13.0% / 20–120 mmol/mol | Should not be used to replace glucose testing in persons with or suspected of having type 1 diabetes, who are pregnant, or a paediatric patient. Should not be used to diagnose diabetes in a range of cases (see package insert). Also not suitable for patients with high blood loss volume, disease affecting erythrocyte cell survival, RF factor | NA | EDTA may cause negative bias. High levels of Haemoglobin F, Haemoglobin S, Haemoglobin C, or other haemoglobin variants may cause interference | Blu-Dock available separately | 18–25°C | <80% | NA | Yes | NA | A1CNow+ System – PTS Diagnostics | 71.69 (INR 6000) | NA | NA | Available | Available INR 6000 | Available | NA | NA | NA | NA | Available | NA | NA | NA | NA | https://www.indiamart.com/proddetail/a1c-now-plus-diabetes-diagnose-system-14855653588.html PTS Diagnostics India Private Limited https://ptsdiagnosticsindia.com/product/a1cnow/ | PT. Mitra Asa Pratama | NA | NA | NA | NA | Professional Medical Science Co., Ltd. | NA | NA | NA | NA | INR 10000 https://www.indiamart.com/proddetail/a1c-now-plus-diabetes-diagnose-system-14855653588.html | NA | NA | NA | NA | NA | NA | NA
11 | HbA1c test | LumiraDx | CE-IVD |  |  |  | IFCC | NGSP |  | Immunoassay |  | Test strips |  | Benchtop |  | Not reported | 0.8 | No | The LumiraDx HbA1c test is an in vitro diagnostic test for the quantitative determination of hemoglobin A1c (IFCC mmol/mol and NGSP %) in human capillary and venous whole blood samples | Professional use only | Capillary fingerstick or venous whole blood | 15 μL | <5 min | IFCC(mmol/mol)/NGSP/DCCT (%) values | Insert test strip, collect sample, prepare sample | Apply sample on the strip and run the analyzer | Analyzer: HbA1c test, User manual, adapter. Test: Test strips, product inserts, lysis devices, RFID tag | NA | NA | NA | 2-30°C | 4.0 –14.0% / 20 - 130 mmol/mol | NA | <2.50% | NA | NA | NA | NA | NA | color touch screen | NA | LumiraDx - HbA1c Test | 1732.75 (INR 145000) | NA | NA | NA | Available INR 145000 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | Nishka medicaids, Ahemdabad Gujarat | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
12 | A1cChek Pro | Wuxi BioHermes | CE-IVD |  |  |  | IFCC | NGSP | Boronate affinity |  |  | Test strips |  | Benchtop |  | 341 × 266 × 234 | 4.4 | No | The A1cChek Pro Glycohemoglobin Analyzer, healthcare professional are able to acquire multiple results only within several minutes to conduct efficient screen of patients at risk of diabetes and make timely medical decision to diabetic patients. | Professional use only in physicians’ office or health clinics | Capillary whole blood, anticoagulated venous whole blood | 3 µL | 5 min | IFCC(mmol/mol)/NGSP/DCCT (%) values | Collect blood samples. Add the blood to the strips | Buffer bottle installation, code chip insertion, ID input, Strip insertion, Blood adding | Analyzer: A1cChek Pro Glycohemoglobin analyzer, power adapter, user manual, quick guide, cleaning and maintenace guide, warranty hangtag Test: Blood collector, Test strip, sampler, code chip, Buffer A, Buffer B, Blood collector. | Cotton swab, alcohol swab, lancet | Level 1 and level 2 | Auto calibration | -20~55℃ | 4.0 –14.0% / 20 - 130 mmol/mol | NA | <3% | NA | USB, RJ45 LAN, RS232-C | 10-40℃ | 30-75% | External power adapter AC input 100-240V~, 1.5A, 50-60Hz; DC output 19.0V, 4.74A | Touch screen | 1000 results | 1. https://www.soymed.com/product/glycohemoglobin-analyzer-2/ 2. HbA1c Analyzer เครื่องตรวจวิเคราะห์ฮีโมโกลบินเอวันซี (biozen.co.th) | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | Available | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | Biozen Co.,Ltd., Biozen 23 Soi Onnuch 62, Onnuch, Suanluang, Bangkok 10250, Thailand https://www.biozen.co.th/14717924/hba1c-analyzer | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
13 | A1c EZ 2.0 | Wuxi BioHermes | CE-IVD |  |  |  | IFCC | NGSP | Boronate affinity |  | Cartridge |  |  | Handheld |  | 123 × 62 × 25 | 0.11 | No | For diagnosis of diabetes, to evaluate the effectiveness of glycaemic control of patients, screening of high-risk group of diabetes | Professional use only in physicians’ office or health clinics | Capillary whole blood, anticoagulated venous whole blood | 3 µL | 6 min | IFCC(mmol/mol)/NGSP/DCCT (%) values | Insert the code chip, insert the test strip. Collect blood with sampler. Add 3 drops of buffer A to test strip. Press the sampler thread onto the test strip. Add 3 drops of buffer B | Read the test result | Analyzer: A1c EZ 2.0, user's manual, operation guide, cleaning and maintenance guide, warranty card. Tests: test strips, buffer A, buffer B, blood sampler, code chip, package insert | NA | None | Not required | At room temperature for 12 months | 4.0–14.0% / 20.2–129.5mmol/mol | NA | <4% | No interference from HbF and haemoglobin variants or unstable glycohemoglobin | Bluetooth, USB | 10–40°C | 30-75% | Battery powered 3V CR2032 | LCD display | NA | https://www.biohermes.com/products/show-26.html | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | Available | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | Biozen Co.,Ltd. Biozen23 Soi Onnuch 62, Onnuch,Suanluang, Bangkok 10250, Thailand https://www.biozen.co.th/14717924/hba1c-analyzer | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
14 | Labona-Check A1c | Ceragem Medisys Inc | CE-IVD | FDA |  |  | IFCC | NGSP | Boronate affinity |  | Cartridge |  |  | Benchtop |  | 178 x 195 x 75 | 0.73 | No | The LabonaCheck™ A1c is intended for the quantitative measurement of glycated hemoglobin in venous whole blood and capillary fingerstick samples. This device is intended for multiple patient, professional use. Measurement of percent glycated hemoglobin (HbA1c) is effective in monitoring long-term glucose control in individuals with diabetes mellitus by using the LabonaCheck™ A1c. Only auto-disabling, single use lancing devices should be used with this system | Professional use | Capillary whole blood | 5 µl | <3 min | IFCC(mmol/mol)/NGSP/DCCT (%) values | Equilibrate reagents to room temperature, collect blood sample | Insert the cartridge into the analyzer and touch analyze button | Analyzer, adapter,printer paper, calibration plate, user manual | NA | NA | Auto calibration | NA | 4.0–15.0% / 20–140 mmol/mol IFCC | NA | 2.8-3.8% | NA | NA | 15–35°C | 15-75% | NA | Touch screen | 300 tests | https://www.tradekey.com/brochure/LabonaCheck-A1c--HbA1c-Analyzer-pbid-188909.html | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | HbA1c Analyzer เครื่องตรวจวิเคราะห์ฮีโมโกลบินเอวันซี (biozen.co.th) | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
15 | Xpress A1c analyzer | Accurex |  | FDA |  |  | IFCC | NGSP | Boronate affinity |  |  | Test strips |  |  | Handheld | 61.6 x 122.9 x 24.5 | 0.116 | No | Accurex Xpress A1c Glycohemoglobin Analysis System that is the only HbA1c analyser approved by US FDA (510k), is used for testing glyochemoglobin concentration as an aid to monitor the risk of developing diabetes & control the status of diabetes. | Professional use | Capillary fingerstick or venous whole blood | 3 µL | 5 min | IFCC(mmol/mol)/NGSP/DCCT (%) values | Equilibrate the reagents to room temperature, collect sample | Insert the test strip, Apply the sample | Analyzer, A1c manual, operational manual, cleaning and maintenance procedure guide; Test kit: calibration chip, Buffer A, Buffer B, Test strips and sampler. | Lancet, wipes | Level 1 and level 2 | Not required | At 2–8°C, 24 months | 4.0 –14.0% / 20 - 130 mmol/mol | NA | <3% | NA | Mini USB | 10-40℃ | 30-75% | 6.0V, AAA Battery x 4 | NA | 1000 tests | Buy Xpress A1c Meter (HbA1c Analyzer) Glycohemoglobin ¦ Accurex | NA | NA | NA | NA | Available INR 65000 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | Buy Xpress A1c Meter (HbA1c Analyzer) Glycohemoglobin ¦ Accurex https://www.hospitalstore.com/accurex-xpress-a1c-handheld-hba1c-analyzer-approved-by-fda-result-in-just-5-minute/ | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | INR 65000, INR 42480 | NA | NA | NA | NA | NA | NA | NA
16 | Eclipse A1c analyzer | Apexbio | CE-IVD |  |  |  | IFCC | NGSP |  | Immunoassay | Cartridge |  |  | Benchtop |  | 248 x 162 x 95 | 1.2 | No | Eclipse A1c offers a fast and easy way to measure HbA1c in a clinic or long-term care facility that prioritizes speed and convenience | Professionals in laboratories, clinics and hospitals | Capillary whole blood, anticoagulated venous whole blood | 1 µL | 5 min | IFCC(mmol/mol)/NGSP/DCCT (%) values | Collect blood samples, insert the capillary holder into the cartridge | Insert the cartridge into the analyzer and close the lid | Analyzer: Eclipse A1c, Power cable with AC adaptor, Air filter, Reagent cartridge | NA | NA | NA | 10 - 40°C | 4.0 –14.0% / 20 - 130 mmol/mol | NA | <4% | NA | USB | 18–35°C | NA | NA | NA | 500 tests | https://www.isotekindo.co.id/product/details/9 | 959.19 (IDR 14950000) | NA | NA | NA | NA | Available | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 1. https://e-katalog.lkpp.go.id/katalog/produk/detail/1374839?lang=id&type=general 2. https://en.indotrading.com/isotekindointertama/alat-cek-hemoglobin-eclipse-a1c-poc-analyzer-s54133-p872936.aspx PT. Isotekindo Intertama Jl. Raya Kebayoran Lama No.309 C, RT.2/RW.1, Grogol Utara, Kec. Kebayoran Lama, Kota Jakarta Selatan, Daerah Khusus Ibukota Jakarta 12210 Jakarta Selatan, DKI Jakarta, Indonesia | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | IDR 14,950,000 , IDR 13,050,000 | NA | NA | NA | NA | NA | NA
17 | SD A1cCare | SD Biosensor | CE-IVD |  |  |  | IFCC | NGSP |  | Immunoassay |  | Test strips |  | Benchtop |  | 163 x 96 x 56 | 0.45 | No | SD A1cCare is an optimized system applicable at the point-of-care testing field. | Professional use | Capillary whole blood, anticoagulated venous whole blood | 5 µL | 3 min | IFCC(mmol/mol)/NGSP/DCCT (%) values | SD A1cCare™ analyzer for HbA1c testing is designed for monitoring fresh capillary or venous whole blood. Do not use the frozen blood or the artificial materials not allowed. | Insert a test panel, Take 5 µL of blood sample, and mix it with buffer solution, Apply the mixed sample to the test panel, check the result after 3 mins. | Analyzer: A1c Care Analyzer, Check strip, DC adapter, User manual, Quick guide. Test kit: Test setx20, Buffer tube x 20, Code-chip, Test kit insert. HbA1c control: Level 1x 10 , Level 2 x 10 | Lancet, wipes | Level 1 and level 2 |  | 1-30°C, 18 months | 4-15% (20-140 mmol/moL) |  | <3% |  |  | 15–32°C |  | 5V/2A AC power adapter or 4 AA alkaline batteries can be used |  | 1000 tests | https://ethitech.co.za/wp-content/uploads/2019/05/SD-A1cCare-User-guide.pdf | NA | NA | NA | Available | Available |  | NA | NA | NA | NA | NA | NA | NA | NA | https://www.komachine.com/en/companies/sd-biosensor/models | https://www.hospitalstore.com/sd-bosensor-a1ccare-rt-pcr-machine-price/ | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA

## Multiple tests
S.no | Product Name | [27 empty] | Technical | [5 empty] | Procedure | [7 empty] | Performance |  |  |  | Analyzer | [5 empty] | Resources | Cost (USD) | Availability | [10 empty] | Supplier | [10 empty] | Cost
|  | Manufacturer | SRA Approvals | [8 empty] | Certifications |  |  |  | Method |  |  |  | Test Design |  |  |  | Device Design |  | Dimensions in mm | Device weight in kg | Additional Test | Use | User | Type of Sample | Volume | Duration | Reporting Units | Sample Prep | Handling | Equipment | Not provided | Control | Calibration | Storage temperature and shel life |  | Reference Range | Limitation | Precision | Interference | Data Export | Operating temperature | Humidity | Power | Display | Memory |  |  | Bangladesh | Bhutan | Republic of Korea | India | Indonesia | Maldives | Myanmar | Nepal | Sri Lanka | Thailand | Timor-Leste | Bangladesh | Bhutan | Republic of Korea | India | Indonesia | Maldives | Myanmar | Nepal | Srilanka | Thailand | Timor-Leste | Bangladesh | Bhutan | Republic of Korea | India | Indonesia | Maldives | Myanmar | Nepal | Sri Lanka | Thailand | Timor-Leste
1 | Afinion 2 | Abbott | CE-IVD | FDA | CLIA waived, Japan | [6 empty] | IFCC | NGSP |  |  | Boronate affinity |  |  |  | Cartridge |  |  |  | Benchtop |  | 200 x 186 x 328 | 3.4 | Lipid panel, CRP, ACR | Afnion HbA1c is an vitro diagnostic test for quantitative determination of glycated hemoglobin (hemoglobin A1c, HbA1c) in human whole blood. This test can be used as an aid in the diagnosis of diabetes and as an aid in identifying patients who may be at risk for developing diabetes. | Professional near patient testing and laboratory | Capillary whole blood, anticoagulated venous whole blood | 1.5 μL | 3 min | IFCC(mmol/mol)/NGSP/DCCT (%) values | Collect the sample with the integrated sampling device. Place the sampling device back in the test cartridge | Place the test cartridge in the analyzer and close the lid | Analyzer:Afinion 2 analyzer, power cable, power supply, 24 VDC, user manual. Test:test cartridges packaged separately in foil pouches with a desiccant bag, package insert. | Standard blood collection equipment | Afinion HbA1c control is a two point control kit with data | Each lot of an Afinion Test includes calibration data stored in the barcode label. The calibration data is read by the integrated camera and used for calculating results. | At 2-8 degree C, 24 months; At 15-25 degree C, 3 months |  | 4.0-15.0%/20-140mmol/mol | Any cause of shortened erythrocyte life span results in a decrease in HbA1c values. Caution should be used when interpreting results from patients with hemolytic anemia or other hemolytic diseases, hmozygous sickle cell trait, pregnancy, blood loss, polycythemia, iron deficiency etc. This test should not be used to diagnose: diabetes during pregnancy, patients with elevated fetal hemoglobin (Hb>10%), patients with a hemoglobinopathy, patients that received a blood transfusion within the past three weeks, patients that have received cancer chemotherapy within the past 3 weeks. | <3% | No significant interference (<7%) was observed up to the following concentrations: Bilirubin conjugated 600 mg/L, Bilirubin unconjugated 600mg/L, Lipids 10g/LRheumatoid factor 780000 IU/L, Total protein 150g/L, Glycated albumin 7.7g/L | Ethernet cable connection to LMIS systems using POCT1-A, HL7, ASTM 1381-85 or ASTM 1394-97 protocols | 18-30 °C | 10-80% (non condensing) | Seperate AC to DC power supply. Double installed | Standard LCD colour display with back light and integrated touch panel | 500 patient tests/500 controls | 1.https://www.skup.org/media/wy0b323c/2021-126-afinion-2-analyzer-hba1c-rapport.pdf 2.https://accessdata.fda.gov/scripts/cdrh/cfdocs/cfRES/res.cfm?id=180494 | 4738.28 https://stat-technologies.com/product/afinion-2-analyzer-pricing-coming-soon/ | NA | NA | NA | Available | Available | Available | Available | NA | NA | Available | NA | NA | NA | NA | 1. https://www.globalpointofcare.abbott/in/en/product-details/afinion-hba1c.html 2. Sai Ram Enterprises 3. Saaskin Corporation Private Limited https://www.saaskin.in/ | "Medquest", Menara Salemba 6th Fl Zone 1, Jl. Salemba Raya No. 5 - 5A, Jakarta Pusat 10440, +62-21 39842961 https://medquest.co.id/products/clinical-chemistry/abbott-afinion-2-analyzer/ | "Medtech Maldives Pvt.Lt", Ma. Fas Eri 3rd Floor, Ameenee Magu, Male, Male 20205, +960 330-8663 https://medtechmaldives.com/products/afinion-2-7799 | https://www.amdmyanmar.com/product/afinion-2/ | NA | NA | https://www.globalpointofcare.abbott/ww/en/global-sites.html | NA | NA | NA | NA | INR 145000 https://www.indiamart.com/proddetail/afinion-2-analyzer-21166615173.html, INR 190000 https://www.indiamart.com/proddetail/hba1c-abbott-afinion-2-20374733412.html | NA | NA | NA | NA | NA | NA | NA
2 | NycoCard Reader II | Abbott | CE-IVD | [8 empty] | IFCC | NGSP |  |  | Boronate affinity | [5 empty] | Reader Pen |  | Benchtop |  | 200 x 170 x 70 | 0.54 | CRP | NycoCard HbA1c is an in vitro diagnostic medical device for quantitative determination of glycated hemoglobin (HbA1c) in human whole blood. For monitoring of long-term metabolic control in persons with diabetes mellitus | For professional near-patient testing and laboratory use | Capillary whole blood, anticoagulated venous whole blood | 5 µL | 3 min | IFCC(mmol/mol)/NGSP/DCCT (%) values | Equilibrate reagents to room temperature. Add sample to reagent tube, mix well and leave the tube for 2–3 minutes. Then add sample to test device and apply washing solution. | Measurement is not automated, it is performed with a pen device | Analyzer: Nyocard Reader II, battery charger, batteries (rechargeable), calibration devices 10 units, pen tips, pen rings, pad. Test: test device, reagent, washing solution | Capillary tube or pipette (5 μL) for sample collection, Capillary tube holder, Pipette (25 μL) and pipette tips | NycoCard HbA1c Control consists of two liquids, each 1,5 ml, Level I+II (normal/abnormal) | Black and white calibration of reader pen. Lot-specific calibration not required | 2–8°C |  | 4.0–15.0% / 20–140 mmol/mol | Hemolysed samples with plasma Hb >3 g/100 mL will interfere with the test system | <5% | The following hemoglobin (Hb) variants have been analysed and found not to affect the NycoCard HbA1c Test result: HbAC, HbAE, HbAF, HbAJ, HbAS and Hb Yamagata. Carbamylated and pre-glycated hemoglobin does not affect the test. Elevated amounts of glucose, bilirubin, lipids and fructosamine were added to blood samples with normal and elevated HbA1c values. No interference was obtained | NA | 15–35°C (The recommended range is 18°C–25°C) | 0–90% (non-condensing) | Rechargeable NiMH batteries only. Size AA 1.2 V. Recharge 8–18 hours | LCD screen 2×16 Characters | 2 KB | https://www.globalpointofcare.abbott/ww/en/product-details/nycocard-reader.html | 774.15 ( INR 65000) | NA | NA | NA | Available | Available | NA | Available | NA | NA | Available | NA | NA | NA | NA | 1."Diagnosol Healthtech ", Jhotwara, Jaipur, Rajasthanhttps://www.diagnosolhealthtech.com/ , 2. "Ank Cares Private Ltd", Ashram Road, Ahmedabad, Gujarathttps://www.ankcares.in/contact-us.html | NA | NA | Advanced Medical & Diagnostics Trading’s , Yangon: 01-571656, 09-7316248 amd@amdmyanmar.com https://www.amdmyanmar.com/product/afinion-2/ | NA | NA | DIAGNOSTICS Abbott Laboratories Ltd., No.1, Q.House Lumpini 30th Floor South Sathorn Road Thungmahamek, Sathorn Bangkok 10120, Thailand Telephone: (66 2) 697 2888 https://www.abbott.com/global-locations-contacts/thailand.html | NA | NA | NA | NA | INR 65000 https://www.indiamart.com/proddetail/abbott-nycocard-reader-ii-2853605205597.html, INR 42000 https://www.indiamart.com/proddetail/abbott-nycocard-reader-24714627530.html | NA | NA | NA | NA | NA | NA | NA
3 | Afias-1 | Boditech Med | CE-IVD | [8 empty] | IFCC | NGSP |  |  |  | Immunoassay |  |  | Cartridge |  |  |  | Benchtop |  | 320 x 204 x 180 | 3.9 | Cardiac, Cancer, Hormone, Infection and others (Vit D, Ferritin, Anti-CCP plus) | AFIAS-1 is a single channel lateral flow immunoassay platform. The disposable cartridge includes all the reagents necessary for the test. | Professional use only (Healthcare professionals or trained medical technologists) | Capillary whole blood, anticoagulated venous whole blood, fingerprick | 10 µL, 30 µL, 50 µL | 3-15min | NGSP (%) | Collect the blood using a C-tip and place it on the Cartridge | Insert the cartridge into the analyzer and press start | Analyzer: Afias-1, AC adaptor and power cable, thermal printer paper roll, Afias OS SD card, operation manual, system check cartridge set | NA | Two controls, QC 1 and QC2 | Two calibrators | 2 – 8℃ , 20 months for reagents , -20 -50 ℃ for equipment |  | 4.0–15.0% / 20–140 mmol/mol for HbA1c | NA | NA | NA | RS 232/USB port/Ethernet/SD card slot LIS/HIS compatible | 15–35°C | 10- 70% | 100 – 240V AC, 50-60Hz | 5 inch touch screen | 5,000 patient results, 5000 QC results, 500 ID Chips, 100 User IDs | https://www.boditech.co.kr/en/support/id/246 |  |  |  | Available | Available |  |  | Available |  |  | Available |  |  |  | “ Boditech.co” 43, Geodudanji 1-gil,Dongnae-myeon, Chuncheon-si, Gangwon-do, Republic of Korea https://www.boditech.co.kr/en/product/instruments/id/4 T. 82-33-243-1400 | 1. https://hospitalstore.in/ 2. https://jaindiagnostic.com/ 3.https://labx.in/product/afias-1/ | "e-katalog " https://e-katalog.lkpp.go.id/katalog/produk/detail/77819305?lang=id&type=general |  | Healthcare Mandalar Pharmaceuticals and Laboratory Company Limited , https://hmpl.com.mm/en/publication/afias-1 Phone number: 959 797 357 000, +959 797 357 007 |  |  | 1. Pulse Science, 28/9 Moo 1 Bangsrimuang, Muang, Nonthaburi 11000, Thailand Phone: +66-2886-7808, https://www.pulsescience.co.th/ivf-pulse/afias-series ​ |  |  |  |  | INR 254999 https://www.hospitalstore.com/boditech-afias-1-immunoassay-analyzer/?srsltid=AfmBOorpBXSXg_urSS2GgtSEt4e4zE7M8il_d-Ztox8SCGxdD-MRYcJE , INR 150000 https://www.indiamart.com/proddetail/immunoassay-analyzer-afias-1-2852998129891.html | IDR 102300.000 https://e-katalog.lkpp.go.id/katalog/produk/detail/77819305?lang=id&type=general
4 | ichroma II | Boditech Med | CE-IVD | [8 empty] | IFCC | NGSP |  |  |  | Immunoassay |  |  | Cartridge |  |  |  | Benchtop |  | 276 x 220 x 91 | 1.3 | Cardiac, Cancer, Hormone, Infection, Autoimmune, Organ function, Other function (Calprotectin, Ferritin, H.pylori SA, Pepsinogen I/II, Vit D, Vit D neo) | ichroma HbA1c is a fluorescence Immunoassay (FIA) for the quantitative determination of HbA1c(Hemoglobin A1c) in human whole blood. It is useful as an aid in management and monitoring of the long- term glycemic status in patients with diabetes mellitus. For in vitro diagnostic use only | Hospital laboratories, physicians’ clinics | Capillary whole blood, anticoagulated venous whole blood | 5 µL | 12 min | NGSP (%) | Equilibrate reagents to room temperature. Transfer 100 μL of hemolysis buffer into detection buffer tube. Draw 5 μL of fingertip blood or tube blood using capillary tube. Put capillary tube into the detection buffer tube. Mix the sample thoroughly. Take out the cartridge from the i-Chamber (incubator) slot. Pipette out 75 μL of the sample mixture and load it into a sample well in the test cartridge. Insert the cartridge into i-Chamber slot. Leave for 12 minutes | Insert the cartridge into the analyzer | Analyzer: ichroma II, AC adaptor and power cable, operation manual, system check cartridge set. Test: cartridge, detection buffer tubes, haemolysis buffer vial, ID chip | Standard blood collection equipment. 5 μL Capillary tube. i-Chamber (incubator) | Boditech HbA1c Control consists of two liquids (level 1 and level 2). Carefully reconstitute each vial of lyophilized liquid with exactly 0.5 mL of sterilized distilled water. | Factory-calibrated Lot-specific calibration | Cartridge and Hemolysis buffer: at 4–30°C, 20 months Detection buffer: at 2–8°C, 20 months |  | 4.0–15% / 20–140 mmol/mol | The test may yield false positive result(s) due to the cross reactions and/or non-specific adhesion of certain sample components to the capture/detector antibodies. The test may yield false negative result. Other factors may interfere with the test and cause erroneous results, such as technical/procedural errors, degradation of the test components/reagents or presence of interfering substances in the test samples. | <5% | No interference from significant variants and interfering substances | USB 4 ports, LAN port, USB OTG port | 15–35°C | 10–70% | DC 12V/5A, AC/DC Adaptor, (AA battery) DC 1.5V X 4ea. Input 100-240V ~ 50/60Hz, 1.8A. Output DC 12V/5A | LCD colour touch screen 153mm x 91mm | 1000 patient tests 500 system check results 1000 controls | ichroma™ II - Boditech Med | 1126.44 (INR 94500) | Available | NA | Available | Available | Available | NA | Available | NA | NA | Available | NA | 1. "Medistore" Suite# F-11, Level# 11 Tropicana Tower, 45 Purana Paltan Dhaka-1000, Bangladesh. Mobile: +880 1405100400 E-mail:contact@medistorebd.comBuy online i-Chroma II Immunoassay Reader in Bangladesh (medistorebd.com) 2. Global Medical Engineering(BD),Ltd. 17/2, Topkhana Road (2nd Floor), Dhaka – 1000, Bangladesh. Mobile: +8801404003501 Phone: +8809678020555 E-mail. info@gmebd.comhttps://gmebd.com/product/hormone-analyzer/ | NA | “ Boditech.co” 43, Geodudanji 1-gil,Dongnae-myeon, Chuncheon-si, Gangwon-do, Republic of Korea https://www.boditech.co.kr/en/product/instruments/id/4 T. 82-33-243-1400 | 1. IChroma II Reader ¦ Kin Diagnostics ¦ Mumbai ¦ Kolkata ¦ Delhi 2. https://quicklab.in/product/i-chroma-ii-immunoassay-analyzer/ 3.https://cpcdiagnostics.in/product/ichroma-ii/ | Zoe Pelita Nusantara Jl. Cikini 2 No. 8b, Menteng. DKI Jakarta, 10330 , (021) 2302118 | NA | https://hmpl.com.mm/en | NA | NA | Home ¦ Medlab Asia & Asia Health ¦ 16-18 July 2025 | NA | BDT 220,000 https://www.medistorebd.com/product/i-chroma-ii-immunoassay-reader/ | NA | NA | INR 90000 https://quicklab.in/product/i-chroma-ii-immunoassay-analyzer/ | NA | NA | NA | NA | NA | NA | NA
5 | Cobas b 101 | Roche Diagnostics | CE-IVD | FDA | [7 empty] | IFCC | NGSP |  |  |  | Immunoassay |  |  | Disc Catridge |  |  |  | Benchtop |  | 320 x 300 x 150 | 4.5 | Lipid panel, CRP | The cobas b 101 is an in vitro diagnostic test system designed to quantitatively determine the % hemoglobin A1c (DCCT/NGSP) and mmol/mol hemoglobin A1c (IFCC) in human capillary and venous whole blood | Professional use in a clinical laboratory or at point of care locations | Capillary whole blood, anticoagulated venous whole blood | 2 µL | 6 min | IFCC(mmol/mol)/NGSP/DCCT (%) values | Equilibrate to room temperature prior to use. Sample is applied directly from the fingerstick or via a pipette when testing venous whole blood. The operator applies the sample to the disc. | Insert the disc into the analyzer | cobas HbA1c Test, power adapter, power cable, optical check disc. Tests: individually-wrapped test discs | Optical check disc, General laboratory equipment (e.g., sample transfer pipette for venous blood or alcohol wipes for the fingerstick), Timer | Built-in quality controls Level 1 and level 2 liquid controls. An optical check is also available | Analyzer is factory-calibrated Calibration information is read from each reagent disc | 2-30 °C, 22 months |  | 4-14% /20-130 mmol/mol | This test should not be used for analyzing samples from patients with conditions causing shortened red blood cell survival, such as hemolytic diseases, homozygous sickle cell trait, pregnancy and significant acute or chronic blood loss. Glycated HbF is not detected by this assay as it does not contain the glycated beta- chain that characterize HbA1c, However, HbF levels (>10%) may result in lower than expected %HbA1c values (DCCT/NGSP). | <4% | Heterozygous presence of the most common hemoglobin variants (HbS, HbC, HbD, HbE, and Hb2) does not interfere. Testing results indicate that there is no significant interference for HbS (≤41%), HbC (36%), HbD (42%), HbE (27%) and HbA2 (6.2%) | USB or direct to PC | 15–32°C | 10-85% (non-condensing) | 100 V to 240 V AC (+/-10%), 50/60 Hz | Integrated touchscreen | 5000 patient tests. 500 controls | https://diagnostics.roche.com/global/en/products/instruments/cobas-b-101-ins-2892.html | 5166.67 https://www.vitalmedicalsupplies.com.au/catalogue-products/diagnostic/eye-ear-nose--throat/parts-and-accessories/cobas-b-101-system/ | Available | NA | Available | Available | Available | NA | NA | NA | NA | NA | NA | Biotech Services ¦ country's leading IVD products provider (biotechservices-bd.com) 57/3 & 57/4, 9th floor, ENA Tower Biruttam kazi Nurruzzaman. Road, Sheikh Rasel Square, Lake Circus, Kalabagan, Dhaka-1205 | NA | 한국로슈 (roche.co.kr) 4F, 22, Teheran-ro 108-gil, Gangnam-gu, SEOUL Seoul 06174, South Korea | Roche Diagnostics India Private Limited, Chakala Road Andheri East, Mumbai, Maharashtra https://diagnostics.roche.com/global/en/contact-forms/contact-us-support.html | e-katalog.lkpp.go.id/katalog.produkctr/getdetailproductcenter?id=1468662 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | IDR 643.000,00 https://ekatalog.lkpp.go.id/katalog.produkctr/getdetailproductcenter?id=1468662 | NA | NA | NA | NA | NA | NA
6 | DCA Vantage | Siemens Healthineers | CE-IVD | FDA | CLIA waived | [6 empty] | IFCC | NGSP |  |  |  | Immunoassay |  |  | Cartridge |  |  |  | Benchtop |  | 287 x 277 x 254 | 3.88 | ACR | The DCA Vantage HbA1c assay provides a convenient, quantitative method for measuring the percent concentration of hemoglobin A1c in blood. The measurement of hemoglobin A1c, recommended for monitoring the long-term care of persons with diabetes | Professional use in a physician’s office or hospital laboratory | Capillary whole blood, anticoagulated venous whole blood |  | 6 min | IFCC(mmol/mol)/NGSP/DCCT (%) values | Equilibrate to room temperature prior to use. Obtain sample using capillary holder device. Add capillary holder device to test cartridge. Scan cartridge bar code. Load and pull foil tab | Testing begins when compartment door is closed | Analyzer: DCA Vantage, cleaning kit, optical test cartridge, power cord, air filter replacement kit, spare fuses, quick reference guides, documentation CD, paper roll, self adhesive label stock. Tests: test cartridges, sample capillary holders, calibration card | Lancing devices, gloves, gauze, bandages, sharps bin, alcohol wipes | DCA Controls Kit contains: vials of A1c Normal control, vials of A1c Abnormal control, vial of reconstitution fluid, eye-dropper caps, control card with normal and abnormal values (doublesided), instructions for use | Factory-calibrated. Lot-specific calibration | At 2–8°C, 24 months |  | 2.5–14.0% / 4–130 mmol/mol | The DCA HbA1c assay gives accurate and precise results over a range of total hemoglobin of 7–24 g/dL. Patients with severe anemias may have hemoglobin concentrations lower than 7 g/dL, and patients with polycythemia may have hemoglobin concentrations above 24 g/dL. Conditions such as hemolytic anemia, polycythemia, homozygous HbS, and HbC, can result in decreased life span of the red blood cells, which causes HbA1c results to be lower than expected. | <3% | At levels of hemoglobin F less than 10%, the DCA system accurately indicates the patient’s glycemic control. However, at very high levels of hemoglobin F (> 10%), the amount of HbA1c is lower than expected. Common HbA1c interferents have been tested. None has been found to have significant effects. | Via USB flash drive to PC or direct to LIS/HIS or data manager | 5–40°C | 15–90% | 100 to 240 VAC; 50/60 Hz | Color touch screen with 1/4 VGA resolution | 4000 tests | https://www.siemens-healthineers.com/en-in/diabetes/diabetes/dca-vantage-analyzer | 1547.51 (INR 129,800) | NA | NA | NA | Available INR 129,800 | Available | NA | NA | NA | NA | Available | NA | NA | NA | NA | https://www.siemens-healthineers.com/en-in/corporate-info | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | INR 1,25,000 | NA | NA | NA | NA | NA | NA | NA
7 | Atellica DCA analyzer | Siemens Healthineers | CE-IVD | FDA | CLIA waived | [6 empty] | IFCC | NGSP |  |  |  | Immunoassay |  |  | Cartridge |  |  |  | Benchtop |  | 158.8 x 152.4 x 221.0(HxWxD) | 2.16 | ACR | Atellica DCA Analyzer is compact, portable, and fully customizable | Professional use | Capillary whole blood, anticoagulated venous whole blood | 1 µL | 4 min | IFCC(mmol/mol)/NGSP/DCCT (%) values | No pretreatment or pipetting required | Take a blood sample and directly dispense it to the HbA1c cartridge. Insert the cartridge into the analyzer and run the test | Analyzer, self-contained immunoassay cartridges, User manual, adaptor, power cables | NA | NA | None required. Calibration data is embedded on a bar code on each cartridge | 2–8°C for 60 days |  | 4.0–14% (20–130 mmol/mol) | NA | Not reported | No clinically significant interferences from common hemoglobin variants for HbA1c measurements. Atellica DCA HbA1c test is not affected by hemolyzed samples and can help reduce preanalytical errors | USB Type C port for patient data export | 15–32°C | 10-90% | 100–240 VAC; 50/60 Hz | Data can be entered via the removable handheld color touchscreen, integrated 1D/2D bar-code scanner, or optional external handheld bar-code scanner | 10,000 tests | https://www.siemens-healthineers.com/en-th/diabetes/diabetes/atellica-dca-analyzer | NA | NA | NA | NA | NA | Available | NA | NA | NA | NA | Available | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | INR 1,29,800 | NA | NA | NA | NA | NA | NA | NA
8 | InnovaStar | DiaSys | CE-IVD | [8 empty] | IFCC | NGSP |  |  |  |  | Enzymatic |  | Cartridge |  |  |  | Benchtop |  | 200 x 170 x 150 | 4 | CRP, glucose/haemoglobin | The InnovaStar is an analyzer for biochemical in-vitro diagnostics. oneHbA1c IS is the diagnostic reagent for quantitative in vitro determination of hemoglobin A1c in whole blood on InnovaStar. | For professional use only | Capillary whole blood, anticoagulated venous whole blood | 10 µL | < 5 min | IFCC(mmol/mol)/NGSP/DCCT (%) values | Bring reagent to room temperature. Make sure that the reagent is at the bottom of the cartridge. Take the patient sample with open-end capillary as described in the user manual. Put the filled capillary in the sample cup. Mix the sample | Place sample cup into the analyzer and start the measurement directly | Analyzer: Innovastar, User manual, adapter | Standard blood collection equipmemt. Sample cups InnovaStar 10/500 (magenta cups) and open-end capillaries (10 μL/ heparinized) | Two level control solutions: TruLab HbA1c liquid L1 TruLab HbA1c liquid L2 | Factory-calibrated Pre-calibrated tests | 2–8°C until expiry date |  | 3.0–14.0% / 9–130 mmol/mol | The limit of detection is 30 mmol/mol HbA1c (4.9% hbA1c DCCT/NGSP). In very rare cases, samples of patients with gammopathy might give falsified results. Heterophile antibodies in patient samples may cause falsified results. | <2% at 6.24% DCCT | The variants AS, AC, AD, AG, DD and elevated A2 showed no significant interferences. The variants AE, AJ, SS, CC, SC, SE, EE, elevated F and elevated A2/F can lead to deviant HbA1c results (> 10% IFCC; > 7% DCCT/NGSP). | RS232 connection; USB connection via adapter | NA | NA | Not reported | 2.8 inch colour touchscreen | 50 results | https://www.diasys-diagnostics.com/products/poct-systems/innovastar/ | NA | NA | NA | NA | Available | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | DiaSys India | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
9 | EVYA | BlueSemi.- an Indian startup |  |  |  | ISO 15197 | [5 empty] | NA | [6 empty] | None (Non-Invasive) | Non invasive health monitoring gadget |  |  |  |  | Handheld | 50 x 100 x 20 | 0.15 | Blood glucose, ECG,BP,Heart rate,oxygen | place the thumb over the sensors on the device and hold it for approximately 60 seconds to obtain results. | Personal health monitoring | Not required | Not required | 1 min | Not reported | No preparation required | Place the thumb over the sensors on the device and hold it for approximately 60 seconds to obtain results. | EVYA, USB - C charging cable | NA | NA | NA | NA |  | NA | NA | Not reported | NA | Via Bluetooth to Mobile phone | NA | NA | Battery | No display | NA | https://www.eyva.io/ | 198.47 (INR 16650) | NA | NA | NA | Available INR 16650 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
10 | Getein 1100 Immunofluorescence Quantitative Analyzer | Getein Biotech Inc. | CE-IVD |  |  | ISO | NMPA |  |  |  |  | IFCC | NGSP |  |  |  | Immunoassay |  |  | Cartridge |  |  |  | Benchtop |  | 261 x 241 x 115 (DxWxH) | 1.9 | Cardiac marker - cTnI, NT-proBNP, D-Dimer, CK-MB, CK-MB/cTnI/Myo, H-FABP,CK-MB/cTnI/H-FABP. Renal function - CysC, mAlb, NGAL, β2-MG. Inflammation - hs-CRP+CRP, PCT, SAA. Hormones- TSH,T3,T4. Bone Metabolism- 25-OH-VD. | Getein1100 Immunofluorescence Quantitative Analyzer is an advanced on-site diagnostic analyzer, intended to deliver accurate testing results in minutes just from 3-4 drops of human specimens. | Professional use in hospital ,clinical laboratries | Capillary whole blood, anticoagulated venous whole blood | 10 µL | 5 min | IFCC(mmol/mol)/NGSP/DCCT (%) values | Collect the blood sample | Sample transfer, Test card insert, click start icon, Result show and print | Analyzer: Getein 1100 instrument, Power adapter, data cable, user manual, printing paper. Test kit: 25 individually packed testing cards in a sealed pouch with desiccant, product specifications, and one SD card, sample diluent, capillary pipet | Lancet, alcohol wipes |  | RFID card calibration | 4- 30°C, 24 months (Test kit) |  | 2.0 –14.0% |  |  |  | PC/LIS/HIS | 10℃ ~ 35℃ | ≤ 70% | 100 - 240V～, 50/60Hz, 60VA | 7 inch touch screen | 10000 data | https://www.getein.com/getein-1100-immunofluorescence-quantitative-analyzer_p24.html |  | Available | NA | NA | Available | NA | NA | NA | NA | NA | NA | NA | 1. Techno Healthcare services https://www.technohealthcare.com.bd/product/37 2. Jist Life care https://jistlifecare.com/product/gp-getein-1100 3. LabPro Bangladesh https://labprobd.com/product/gp-getein-1100-immunofluorescence/ | NA | NA | 1."Diagnosol Healthtech ", Jhotwara, Jaipur, Rajasthanhttps://www.diagnosolhealthtech.com/ 2.https://www.cosmic.net.in/product/getein-1100-immunofluorescence-quantitative-analyzer/ | NA | NA | NA | NA | NA | NA | NA | BDT 45000 https://www.technohealthcare.com.bd/product/37 | NA | NA | INR 52000 https://www.diagnosolhealthtech.com/, INR 125000 https://www.cosmic.net.in/product/getein-1100-immunofluorescence-quantitative-analyzer/ | NA | NA | NA | NA | NA | NA | NA
11 | GETEIN 1160 IMMUNOFLUORESCENES QUANTITATIVE ANALYSER | Getein Biotech Inc. | CE-IVD | [8 empty] | IFCC | NGSP |  |  |  | Immunoassay |  |  | Cartridge |  |  |  | Benchtop |  | 299 × 276 × 152 (W × D × H) | 4 | Renal Function- CysC, mAlb, NGAL, β2-MG, Cardiac markers - cTnI, TnT, BNP, NT-proBNP, H-FABP, CK-MB, CK-MB/cTnI, CK-MB/cTnI/Myo, CK-MB/cTnI/H-FABP. Infectious diseases - Anti-HCV, Anti-TP, Anti-HIV, SARS-CoV-2 Neutralizing Antibody, HBsAg, Anti-HBs, SARS-CoV-2 Antigen, H. pylori, Influenza A/B, Dengue NS1 Ag. Thyroid function - TSH, T3, T4, fT3, fT4. Inflammation - hs-CRP+CRP, PCT, SAA, IL-6. Reproduction - HCG+β, FSH, LH, PRL, AMH, Prog, Testosterone, E2. Specific Protein and Rheumatism - RF, ASO, Anti-CCP. Others- Ferritin, Total IgE, PG I/PG II | The Getein1160 Immunofluorescence Quantitative Analyzer is used for the detection and quantification of various biomarkers in biological samples. | Professional use in hospital ,clinical laboratries | Capillary whole blood, anticoagulated venous whole blood | 10 ~ 200 μL depending on the assay | 3-10 min | IFCC(mmol/mol)/NGSP/DCCT (%) values | Collect the sample using appropriate device | Insert the test cartridge into the analyzer | Analyzer, test cards, USB, power supply, user manual. Test kit:Sealed cartridge with 24/48 Getein HbA1c test cards, User manual | Standard blood collection equipment, sample diluent, box with pipette tips, Mixing plate | NA | Factory calibrated | Analyzer -40℃ to +55℃ , 4-30°C, 24 months |  | 2.0 –14.0% | The analyzer is intended solely for in vitro diagnostic analysis of human whole blood, plasma, serum, capillary blood, and urine samples. It should not be used for any other purposes and users must be trained on the operations and safety precautions | <2% | NA | USB | 10℃ ~ 35℃ | ≤ 70% | 100 - 240V～, 50/60Hz, 60VA | LED touch screen | 100000 | Chinese Getein 1160 Immunofluorescence Quantitative Analyzer Suppliers ¦ Getein | 14950 https://arg.health/product/gp-1160/ | Available | NA | NA | Available INR 125000 | NA | NA | NA | NA | NA | NA | NA | Biotech International-Leading Medical Equipment supplier in Bangladesh (bio-techbd.com) | NA | NA | 1. Global biomedicals, Rajasthan 2. alphalifecare.co.in 3.https://www.saaskin.in/ | NA | NA | NA | NA | NA | NA | NA | BDT 65000 https://www.medistorebd.com/product/getein-1160-poct-hormone-analyzer/ | NA | NA | INR 125000 | NA | NA | NA | NA | NA | NA | NA
12 | Allegro | Nova Biomedical | CE-IVD | [8 empty] | IFCC | NGSP |  |  |  | Immunoassay |  |  | Cartridge |  |  |  | Benchtop |  | 457 x 356 x 381 | 10.43 | Lipid panel,CRP, blood glucose, blood creatiine, PT/INR, urine creatinine, urine albumin | Quantitative measurement of HbA1c in whole blood | Professional use | Capillary whole blood | 1.5 μL | <7 min | IFCC(mmol/mol)/NGSP/DCCT (%) values | If refrigerated, equilibrate to room temperature before use. Scan cartridge. Perform fingerstick. Remove capillary sampler from cartridge. Touch sampler to blood. Replace capillary sampler in cartridge | Insert cartridge into analyzer and press "analyze" | Analyzer: Allegro, power cable, companion StatStrip meter Tests: cartridge with capiallary sampler | Lancet, wipes | Level 1 and level 2 | Factory-calibrated Lot-specific calibration | At 2–8°C, 18 months |  | 4.0–14.0% / 20–130 mmol/mol | Diluted samples cannot be analyzed. Do not analyze hemolyzed coagulated blood samples | <4.5% | No significant interference ( | USB port, ASTM Protocol, via serial RS232 TCP/IP, POCT1-A2, HL7 | 18–25°C | 10–80% | 90–264 V AC, 50/60 Hz | Yes | Stores and presents trend data for patients over the last 9 visits | A Fast, Allegro® Simple, Capillary Blood Analyzer (novabiomedical.com) | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | Allegro® ¦ Nova Biomedical Corporation (omnia-health.com) | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
13 | A1 care analyzer | i-SENS | CE-IVD | [8 empty] | IFCC | NGSP |  |  |  |  | Enzymatic |  | Cartridge |  |  |  | Benchtop |  | 290 x 250 x 130 | 3.8 | ACR | The A1Care Analyzer is a bench top, point-of-care (POC) in vitrodiagnostic (IVD) device for the measurement of HbA1c (%, mmol/mol) in whole blood samples | Professional use | Capillary whole blood, anticoagulated venous whole blood | 2.5 µl | 4 min 20 seC | IFCC(mmol/mol)/NGSP/DCCT (%) values | 20 min warm up for cartridges if stored in the fridge. The sample collection device is in the cartridge. Remove it. Collect blood from finger | Insert the test cartridge into the analyzer. Press the sample collector into the cartridge | Analyzer: A1Care analyzer, optical check cartridge, quick guide, user manual, power cord. Tests: cartridge with sample collector | Lancet, wipes | A1Care HbA1c Control Kit consists of tqo liquids: level 1 (normal) and level 2 (abnormal) | Factory-calibrated Lot-specific calibration | At 1–30°C, 24 months |  | 4.0–15.0% / 20–140 mmol/mol IFCC | Total haemoglobin must be within 7–24 g/dL, very high levels of haemoglobin F (>10%), haemolytic anaemia, polycythaemia, homozygous HbS and HbS will cause lower than expected HbA1c | <3% | A range of substances tested | printer or via USB device, LAN port for internet cable connectivity | 10–32°C | 10–90% | AC mains supply (Power Input 100–240VAC 50/60Hz, 100VA | 7 inch LCD colour touchscreen | 5000 patient tests 5000 controls | https://i-sens.com/point-of-care-testing/ | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
14 | AS100 Alere Afinion Analyzer | Alare Technology | CE-IVD | FDA | CLIA waived | [6 empty] | IFCC | NGSP |  |  | Boronate affinity |  |  |  | Cartridge |  |  |  | Benchtop |  | 320 x 170 x 170 | 5 | Lipid panel,CRP, ACR | The Alere Afinion™ AS100 Analyzer is a compact, rapid, multi-assay analyzer that provides valuable near patient testing at the point-of-care. | Professional use | Capillary whole blood, anticoagulated venous whole blood | 1.5 μL | 3 min | IFCC(mmol/mol)/NGSP/DCCT (%) values | Collect the sample using appropriate device | Insert cartridge into analyzer | Analyzer, adapter,user manual | NA | NA | Not required | NA |  | 4.0–15.0% / 20–140 mmol/mol IFCC | The device cannot analyze hemolyzed or coagulated samples. If such samples are tested, the analyzer will display an information code and no result will be obtained. | <3% | It can handle HbF concentrations up to 10.4% without significant interference, higher concentrations may lead to negative interference in the results | USB | 15–32°C | 10–90% (non-condensing) | 100 to 240 VAC; 50/60 Hz | Touch screen | 1000 patient results | Alere Afinion AS100 Analyzer Product Demo Library ¦ Abbott Point of Care ¦ Abbott Point of Care (globalpointofcare.abbott) | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | HbA1c Analyzer เครื่องตรวจวิเคราะห์ฮีโมโกลบินเอวันซี (biozen.co.th) | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
15 | SD multicare analyzer | SD Biosensor, Korea | CE-IVD | FDA | [7 empty] | IFCC | NGSP | JDS |  |  | Immunoassay |  |  |  | Test strips |  |  | Benchtop |  | 163 x 96 x 56 | 0.45 | Lipid Panel, CRP, U-Albumin | MultiCare™ System is a portable reflectometry immunoassay instrument for the quantitative measurement of HbA1c. It is a convenient in-vitro diagnostic. This system is for in-vitro diagnostic use and professional use in hospital and clinic. | Professional use | Capillary whole blood, anticoagulated venous whole blood | 5 µL | 3-5 min | IFCC(mmol/mol)/NGSP/DCCT (%) values | Insert a test panel, apply sample to the test panel , check result after 3 mins | Insert the test panel into the analyzer | Analyzer, adapter, SD HbA1c check strip, user manual, quick guide | NA | NA | Easy calibration by cartridges | 0 °C - 32°C, 18 months |  | 4.0–15.0% / 20–140 mmol/mol IFCC | NA | ≤3% | NA | Mini USB, Bluetooth(optional) | 15–32°C | 10–90% | AC100~240V, 50/60Hz | LCD display | 1000 patient data | Diagnostic Accuracy of Point-of-Care Tests Measuring Glycosylated Haemoglobin (HbA1c) for Glycemic Control: A Field Study in India - PMC (nih.gov) https://www.sdbiosensor.com/link/catalogue/2021%20All%20Product%20Catalog_20210601_low.pdf | 850 (INR 50000) | NA | NA | Available | Available (INR 50000) | Available | NA | NA | NA | NA | NA | NA | NA | NA | SD BIOSENSOR, INC. (Headquarter) Address C-4&5 Floor, 16, Deogyeong-daero 1556beon-gil, Yeongtong-gu, Suwon-si, Gyeonggi-do, 16690, REPUBLIC of KOREA Tel +82-31-300-0400 E-mail sales@sdbiosensor.com Website www.sdbiosensor.com | SD BIOSENSOR (HEALTHCARE PVT. Ltd.) Address SCO-34, Sector-15, Part-II, Gurgaon, Haryana - 122001, India, INDIA Tel +91-124-4540908/Toll Free No. E-mail india@sdbiosensor.com Website www.sdbiosensor.co.in | PT. STANDARD BIOSENSOR INDONESIA Address GD.GL’avenue office tower 21 floor, Unit C, Jalan Raya Pasar Minggu Kav.16. Jakarta Selatan - 12780, Tel +62-21-8066-7161 E-mail purbayu.mulyanto@sdbiosensor.co.id Website www.sdbiosensor.co.id | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
16 | icare - 2100 | Sinocare | CE-IVD | [8 empty] | IFCC | NGSP |  |  | Boric acid affinity |  |  |  | Cartridge |  |  |  | Benchtop |  | 308.6 x 445.0 x 293.0 | <15 | creatinine, ALT, for AST, INR, glucose, lipid, CRP, hemoglobin, cystatin C, HbA1c, total bilirubin, urea, LDH | The iCare-2100 is an innovative portable and automatic multi-function analyzer with accuracy comparable to laboratory results, efficiency within 3-5 minutes, and easy operation only requiring three steps. | Professional use in hospital ,clinical laboratries | Capillary whole blood, anticoagulated venous whole blood | 5 µL for HbA1c | 3 min for HbA1C | IFCC(mmol/mol)/NGSP/DCCT (%) values | Equilibrate the reagent cartridges to room temperature | Insert the reagent cartridge into the analyzer, Insert the sample to the cartridge and press start button | Analyzer, User manual Test kit: Reagent cartridge, Quality control materials, disposable tips, and package insert. | NA | NA | Not required | -20~55℃ for analyzer, 2-8°C for reagent kit |  | 4.0–15.0% / 20–140 mmol/mol IFCC | NA | NA | NA | WiFi, RS232, RJ45, USB | NA | NA | 100 - 240V, Rated power :160 W, Frequency: 50/60 V | 9.7 inch touch screen | 80000 test results, 10000 QC results | POC diabetes analyzer - iCARE-2100 - Sinocare - multiparametric / coagulation / for thrombosis (medicalexpo.com) | 5,597.70(INR 470000) | NA | NA | NA | Available | Available | NA | NA | NA | NA | NA | NA | NA | NA | NA | 1. Biochemistry Analyzer Reagents and Rapid Card Manufacturer ¦ Raj Biosis Private Limited, Jaipur 2.Ank Cares Pvt Ltd 3. https://ptsdiagnosticsindia.com/product/icare2100/ | https://www.isotekindo.co.id/product/details/222 Jl. Raya Kebayoran Lama No.309 C Grogol Utara, Kebayoran Lama Jakarta Selatan 12210 | NA | NA | NA | NA | NA | NA | NA | NA | NA | INR 470000, INR 280000 | NA | NA | NA | NA | NA | NA | NA
17 | QDx Instacheck | DiaSys India | CE-IVD | [8 empty] | IFCC | NGSP |  |  |  | Immunoassay (Fluorescent) |  |  | Cartridge |  |  |  | Benchtop |  | 250 x 185 x 80 (LXWXH) | 1.2 | Vitamin D, Ferritin,Microalbumin, PCT, CRP, Prolactin, Troponin-I, hsCRP, TSH, T3,T4, FSH, LH, Total βHCG | QDx Instacheck™ the most powerful POC diagnostic system that allows patients to have an easy access to the medical treatment, not only in major hospitals ICU’s and NICU’s or laboratories but also in doctor’s clinics and even near the site of patient care. | Professional use | Capillary whole blood | 5 µL | 12 min | IFCC(mmol/mol)/NGSP/DCCT (%) values | Collect sample using a reagent pack | Insert cartridge into the analyzer, Insert reagent pack into the test cartridge and close the lid | Analyzer: DxA1c analyzer, adapter, user manualTest kit: Quick guide, cartridge | Alcohol swab, Lancet | NA | NA | NA |  | 4.0–15.0% / 20–140 mmol/mol IFCC | NA | NA | NA | Mini USB | 15–35°C | < 70% | 100 - 240V AC, 50/60 Hz, 3-5A | LCD display | NA | https://www.diasys.in/fileadmin/Diasys_India/QDx_Flyer/Instacheck_brochure_6_page_new_branding_single_page.pdf?_=1562911496 | 1131.45 (INR 95000) | NA | NA | NA | Available | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 1. "SuppliMed International" Address: B.No-352, Mahtuji Nagar, behind BLOCK-A, near Moze College, Balewadi, Pune, Maharashtra 411045 Phone: 097622 10484 2. https://mkp.gem.gov.in/immunoassay-analyzers/immunoassay-analyzer/p-5116877-94654472421-cat.html#variant_id=5116877-94654472421 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | INR 95000, INR 135000 | NA | NA | NA | NA | NA | NA | NA
18 | Palm Reader | Hipro | CE-IVD |  |  | ISO13485 | [5 empty] | NA |  |  |  |  | Immunoassay (Fluorescent) | [5 empty] | Cassette |  | Handheld | 172 × 80 × 56 | ≤0.22 | TT3, TT4, FT3, FT4, TSH, PSA, IL-6, CRP, HCG, ST2, D-Dimer, PRL, PROG, LH, E2, FSH, AMH, VB12, Folic Acid | Hipro Palm Reader is an immunofluorescence analyzer designed for point-of-care testing | Professional use | Capillary whole blood | 4 µL | 2.5 min | IFCC(mmol/mol)/NGSP/DCCT (%) values | Mix sample and buffer, add to the sample well of the cassette | Insert the cassette into the analyzer, result will be dispalyed | Palm F analyzer, Operating manual of the analyzer and reagent, Test cassette, Buffer, Pipette, and Timer. | NA | NA | NA | NA |  | NA | NA | NA | NA | USB,Bluetooth | 10–30°C | 20-80% | ≤ 1.5 VA | NA | NA | https://www.hipro.us/uploads/file/palm-reader-immunoassay-analyzer-brochure.pdf | 133.61 (BDT 16000) | Available | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | https://www.medistorebd.com/product/hipro-palm-reader-poct-analyzer/ "Medistore" Suite# F-11, Level# 11 Tropicana Tower 45 Purana Paltan Dhaka-1000, Bangladesh. Mobile: +880 1405100400 E-mail:contact@medistorebd.com | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | BDT 16000 https://www.medistorebd.com/product/hipro-palm-reader-poct-analyzer/ | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
19 | FA-160 Fluoroscent immunoassay | Lifotronic | CE-IVD |  |  | ISO13485 | [5 empty] | IFCC | NGSP |  |  |  | Immunoassay (Fluorescent) |  |  | Cartridge |  |  |  | Benchtop |  | 270 x 200 x 100 | 2.7 | D-Dimer,CK-MB,Cardiac Troponin, Hs-CRP, HbA1c, HCG, PCT, mAlbumin, NT- proBNP, COVID-19 IgG, Myoglobin, T3,T4, Cystatin C, CRP, TSH, H-FABP | FA-160 is an immnofluorescence-based diagnostic system with TRFIA (Time resolve fluorescent Immunoassay) detector to measure cardiac diseases, infectious diseases and other major life-style related diseases through testing concentration of various analytes in blood samples. | Professional use | Serum, Plasma, Whole Blood | NA | 3 - 15 mins / test | IFCC(mmol/mol)/NGSP/DCCT (%) values | No sample pre-treatment required | Insert the test strip into the incubation channel, after incubation put the test strip into the test channel, samples tested and results printed | Analyzer, Adapter, Test cartridge, ID chip, User manual, sample diluent | NA | NA | NA | 2-8°C for 12 months |  | 4.0–16.0% / 20–151 mmol/mol IFCC for HbA1c | NA | ≤ 15 % | NA | USB , RS -232, LIS Compatible | 10–30°C | <80% | AC 100V-240V, 50/60Hz, ≤ 50VA | TFT True Color LCD Touch Screen | 1000 tests results | https://fybreeds.com/userfiles/files/FA-160%C2%A0.pdf | NA | Available | NA | NA | Available | Available | NA | NA | NA | NA | Available | NA | 1. Lifotronic- FA-160 Immunofluorescence Analyzer – Health Aid BD House – 30, Road – 03, Sector – 05, Uttara Model Town, Dhaka -1230,Bangladesh +88 01816-202314 2. https://www.medistorebd.com/product-category/hormone-analyzer/page/2/ "Medistore" Suite# F-11, Level# 11 Tropicana Tower 45 Purana Paltan Dhaka-1000, Bangladesh. Mobile: +880 1405100400 E-mail:contact@medistorebd.com | NA | NA | https://www.immunoshop.com/, Belapur, Navi Mumbai, Thane, Maharashtra | https://e-katalog.lkpp.go.id/katalog/produk/detail/76999143 | NA | NA | NA | NA | Home ¦ Medlab Asia & Asia Health ¦ 16-18 July 2025 | NA | BDT 65000 https://healthaid.com.bd/product/lifotronic-fa-160-immunofluorescence-analyzer/ | NA | NA | INR 75000 https://www.indiamart.com/proddetail/fa-160-immunofluorescence-analyzer-22812247797.html?pos=2&DualProdscaps | IDR 29,879,407 | NA | NA | NA | NA | NA | NA
20 | Finecare™ FIA meter plus | Wondfo | CE-IVD |  |  |  | NMPA |  |  |  |  | IFCC | NGSP |  |  |  | Immunoassay | [5 empty] | Cassette | Benchtop |  | 270 x 238 x 146 | 2.8 | 1. CVD- NT-proBNP, cTnI, Myo, CK-MB, H-FABP, 3 in 1 (cTnI/Myo/ CKMB), 2 in 1 (cTnI/NT-ProBNP), BNP 2. Infection - PCT, CRP (hsCRP+CRP), 2 in 1(CRP/PCT), SAA 3. Tumor- AFP, PSA, CEA, fPSA 4. Diabetes and Renal Injury - HbA1c,MAU,Cys C,NGAL,β2-MG 5. Hormone - T3, T4, TSH, Vitamin-D ,Cortisol, Estradiol, Testosterone, fT4 6. Fertility - β-HCG, LH, FSH, PRL,Prog 7. Coagulation - D-Dimer | Finecare™ FIA Meter Plus is a fluorescence immunochromatographic analysing system with internal temperature control, which can help diagnose conditions such as infection, diabetes, cardiovascular diseases, renal injury and cancers, etc. | Professional use at Laboratory, ER, ICU, Clinics, GP office, Ambulance, etc | Capillary whole blood | 10 µL for HbA1c | 3-15 mins | IFCC(mmol/mol)/NGSP/DCCT (%) values | If stored in a refrigerator the buffer should be balanced to room temperature, Mix the sample with buffer, Apply sample onto the catridge | Insert the ID chip into the analyzer, Insert the cartridge and press test | Analyzer: Finecare FIA meter plus, List of contents, Packing list, Power cord, Quick operation instruction for use, printer paper, power adapter, CD disk, Ethernet cable, Operational manual Test kit: Reagent packaging, ID chip, Pipette tips, cassette, buffer, blood sampler quick operation instruction for use. | NA | NA | Self calibration | 2–30°C |  | 4.0 –14.5% / 20 - 135 mmol/mol | NA | NA | NA | LIS/HIS connection | Inner temperature control | NA | 100-240V AC, 50-60 Hz | 8" LED touch screen | 30000 tests results | https://pdf.medicalexpo.com/pdf/mira-lab/finecare-fia-meter-plus/118462-184626-_3.html | NA | Available | NA | NA | Available | NA | NA | Available | NA | NA | NA | NA | https://www.medistorebd.com/product/fine-care-analyzer/ | NA | NA | 1. Star Biomed Private limited, New Rohtak Road, Karol Bagh New Delhi - 110005, India 2. https://www.hospitalstore.com/wondfo-finecare-fia-meter-plus-immunoassay-analyzer/ | NA | NA | https://biomedmyanmar.com/products/11/wondfo | NA | NA | NA | NA | BDT 150000 | NA | NA | 1. INR 115000 https://www.indiamart.com/proddetail/finecare-fia-meter-plus-21385788433.html?pos=5&DualProdscaps 2. INR 79500 https://www.hospitalstore.com/wondfo-finecare-fia-meter-plus-immunoassay-analyzer/ | NA | NA | NA | NA | NA | NA | NA
21 | Aina HbA1c analyzer | Janacare | CE-IVD |  |  | ISO 13485 |  |  | HSA |  |  | NA |  |  |  | Boronate affinity |  |  |  |  | Test strips |  |  |  | Handheld | 79.8 x 36.5 x 16.0 | 0.3 | Lipids, Blood glucose and Hemoglobin | The Aina Blood Monitoring System is a high quality and versatile multi-parameter in-vitro diagnostic platform | Professional use | Capillary whole blood, anticoagulated venous whole blood | 5 µL | 3 min | IFCC(mmol/mol)/NGSP/DCCT (%) values | Collect the blood sample using the capillary holder, drop the collected sample into the blue reagent vial, shake well | Insert the test strip into the device, add the sample onto the strip, add wash buffer, and insert the strip back to the analyzer | Aina device, Battery, Instruction booklet, Pipette, Reagent, wash buffer, Test strips, and capillary holder. | NA | NA | Auto calibration | At 2–8°C until use |  | 4.0–15.0% / 20–140 mmol/mol IFCC | NA | <3% | NA | NA | 18–40°C for HbA1c | NA | NA | NA | NA | https://www.siemens-healthineers.com/en-in/point-of-care-testing/aina | NA | NA | NA | NA | Available | NA | NA | Available | NA | NA | NA | NA | NA | NA | NA | 1. "Vishal Enterprises", Nerul, Navi Mumbai, Thane, Maharashtra 2. Laxmi Sagar, Bhubaneswar, Khordha, Odisha https://www.anandkrishnahospicare.com/ 08047645015 | NA | NA | DKSH (MYANMAR) LIMITED | NA | NA | NA | NA | NA | NA | NA | INR 30000 | NA | NA | NA | NA | NA | NA | NA
22 | Smart 700/340 | Diazyme | CE-IVD | FDA | [5 empty] | KGMP |  | IFCC | NGSP |  |  |  |  | Enzymatic |  | Cartridge |  |  |  | Benchtop |  | 240 x 145 x 140 | 3.4 kg 7.7 lbs | Homocysteine, CRP (c-reactive protein), hs-CRP (high sensitive c-reactive protein), Lipoprotein(a), HbA1c, Microalbumin, D-Dimer, Ferritin, INR (PT), Cystine C (GFR), Troponin | SMART HbA1c Hemoglobin A1c reagents are intended for use with the SMART analyzer for the quantitative determination of stable HbA1c in human capillary and venous whole blood samples. | Professional use | Capillary whole blood, anticoagulated venous whole blood | 20 µL | 8 min | IFCC(mmol/mol)/NGSP/DCCT (%) values | Equilibrate reagents to room temperature. Collect sample | Insert the RFID card into the analyzer, Insert the cartridge, press start |  |  |  |  | At 2-8°C |  | 4.0 –14.0% / 20 - 130 mmol/mol | The linearity of the assay is up to 12% HbA1c. Samples with values above 12% should not be diluted and retested. The assay is formulated for use with human whole blood samples. Total hemoglobin in the sample should be in the range: 9-21 g/dL | <5% |  | USB |  |  | 5 V DC | 6” touch color LCD |  | https://weldonbiotech.com/wp-content/uploads/2018/05/SMART-700-340-Flyer.pdf | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
23 | Epithod 616 | DxGen | CE-IVD | [8 empty] | IFCC | NGSP |  |  | Boronate affinity |  |  |  | Cartridge |  |  |  | Benchtop |  | 14.5 x 21.0 x 24.0 (WXLXH) | 1.8 | TSH, IgG, IgM, CRP, Hemoglobin, D-dimer, hs-CRP, microalbumin, HbA1c, glycated albumin | Epithod 616 is semi-automatic analyzer, quantitatively measures diabetes, infectious and inflammation markers on an easy-to-use platform. | Professional use at laboratory, hospital and POC setting | Capillary whole blood, anticoagulated venous whole blood | 5 µL | < 3 min | IFCC(mmol/mol)/NGSP/DCCT (%) values | Equilibrate reagents to room temperature.collect the blood,drop R1 and W1 | Insert the cartridge into the analyzer | Analyzer, User manual, adapter, supervisor card, printer paper roll, pipetting guide |  |  | Smart NFC card | -20 °C ~ 60 °C |  | 3.0–15.0% / 9–140 mmol/mol | NA | <3% | NA | USB, Bluetooth, Wi-Fi | 15–35°C | 15-75% | DC 12V adpater | 7 inch LCD colour touchscreen | 20000 result memory | https://www.dxgenco.com/ep616 | 600-1200 https://www.tradeindia.com/products/diabetes-test-epithod-616-5262068.html | NA | NA | Available | Available | NA | NA | NA | NA | NA | NA | NA | NA | NA | 1.https://dxgen.vbweb.kr/epithod616 2.https://www.tradekorea.com/product/detail/P794015/Epithod-616-Analyzer.html | https://www.hospitalstore.com/dxgen-epithod616-hba1c-analyzer-hplc/?srsltid=AfmBOoouKWvlN1YtnVEqXjdD6tQR6E3nFB3NmetMGfhnaEDvrrMoN51F | NA | NA | NA | NA | NA | Home ¦ Medlab Asia & Asia Health ¦ 16-18 July 2025 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
24 | Epithod 414 AutoDx | DxGen | CE-IVD | [8 empty] | IFCC | NGSP |  |  | Boronate affinity |  |  |  | Cartridge |  |  |  | Benchtop |  | 19.0 x 27.0 x 28.0 (WXLXLH) | 3.6 | CRP, microalbumin, hemoglobin | Epithod®AutoDx, Fully-automatic analyzer, measures diabetes, infectious, and inflammation parameters quantitatively on an easy-to-use platform. It is suitable for physician office, laboratory, and POC setting. | Physician offices, laboratories and POC settings |  | 5 µL | 3 min | IFCC(mmol/mol)/NGSP/DCCT (%) values | Equilibrate reagnts to room temperature, collect blood using a squeeze capillary | Insert the Item key into the device, Insert the catridge into the device | Analyzer, User manual, adapter, supervisor card, printer paper roll, pipetting guide. Test kit: Test cartridge, Reagent 1, Washing solution, Item key | Pipette tip, Squeeze capillary, Safety Lancet |  | Smart NFC card | -20 °C ~ 60 °C |  | 3.0–15.0% / 9–140 mmol/mol | NA | <3% | NA | Ethernet , Wi-Fi, USB | 15–35°C | 15-75% | DC 12V adpater | 7 inch LCD colour touchscreen | 20000 result memory | https://www.dxgenco.com/116 | NA | NA | NA | Available | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
25 | FA 50 | Genrui Biotech | CE-IVD | FDA |  |  | NMPA | [5 empty] | NGSP |  |  |  | Immunoassay (Fluorescent) | [5 empty] | Cassette | Benchtop |  | 300 x 270 x 160 (LxWxH) | 4 | PCT, hs-CRP, NT-proBNP, cTnl,Myo,CK-MB, cTnl/CK-MB/Myo, D-Dimer, HbA1c, T3, T4, TSH, Cys-C, MALB | FA50 is a simple, innovative and accurate POCT Immunofluorescence analyzer | clinic, doctor's office,ICU and CCU | Whole blood, serum, plasma | 10 µL (for HbA1c) | 10 min | NGSP (%) | Equilibrate the reagents to room temperature, Mix the blood sample with diluent | Magnetic card calibration, pipette sample, Insert cassette | Analyzer, user manual Test kit: Cassette, Magnetic card, Diluent |  |  | Magnetic card auto-calibration | 10°C - 30°C (equipment) |  | 4.0 –14.0% / 20 - 130 mmol/mol | NA | NA | NA | 2 USB ports, LIS | 10–30°C |  | 100-240 AC, 50/60Hz | 5.0-inch touch screen | 50000 results | https://www.genrui-bio.com/products/quantitative-immunoassay-analyzer-fa50.html | NA | NA | NA | NA | Available | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | Geenius Healthcare https://www.indiamart.com/geenius-healthcare/?pid=2853964147712&c_id=567&mid=145262&pn=Genrui%20Biotech%20Fa-50%20Immunoassay%20POC | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | INR 80000 | NA | NA | NA | NA | NA | NA | NA
26 | FIA 8000 | Getein Biotech Inc. | CE-IVD | [8 empty] | IFCC | NGSP |  |  |  | Immunoassay | [5 empty] | Cassette | Benchtop |  | 250 x 250 x 120 (LxWxH) | 1.8 | cTnl, NT-proBNP, NT-proBNP/cTnl, CK-MB/cTnl/Myo, PCT, hs-CRP, D-dimer, HbA1c, mAIb, CysC, NGAL,HCG, β2-MG | FIA8000 Quantitative Immunoassay Analyzer is a medical diagnostic instrument which can be used to measure a variety of biomarkers in human whole blood, serum, plasma or other specimens. | Hospital, Clinic, Laboratory | Whole Blood, Serum, Plasma, Urine | 10 µL | 3 min | IFCC(mmol/mol)/NGSP/DCCT (%) values | Collect blood samples | Insert the QC card, Insert the cassette with the sample into the analyzer, and click measure | Analyzer, User manual, adapter, data cable, QC card, printing paper. Test kit: Cassette, Diluent. |  | Quick quality control with QC card |  | At 2-8°C |  | 2.0 –14.0% | NA | 1% | NA | PC, LIS, HIS connectivity | 15–35°C | 10–85% | AC 100V-240V, 50/60Hz | 5.6 inch touch screen | 10000 results | https://www.getein.com/fia-8000-quantitative-lmmunoassay-analyzer_p31.html | NA | NA | NA | NA | Available | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | Global Biomedicals Inc. https://www.globalbiomedicals.com/ | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
27 | FIA 8600 | Getein Biotech Inc. | CE-IVD |  |  | ISO | NMPA | FSC |  |  |  | IFCC | NGSP |  |  |  | Immunoassay | [6 empty] | Benchtop |  | 220 x 230 x 98 (DxWxH) | 1.5 | Renal function - CysC, mAlb, NGAL, β2-MG; Diabetes- HbA1c; Cardiac Marker - cTnI, NT-proBNP, NT-proBNP/cTnI, CK-MB, CK-MB/cTnI/Myo, CK-MB/cTnI; Coagulation marker - D-dimer; Inflammation - hs-CRP+ CRP, PCT; Thyroid- TSH; Reproductive/Fertility - HCG+β | FIA 8600 Quantitative Immunoassay Analyzer is an analyzer for the processing and analysis of Getein test kits including markers for cardiovascular diseases, renal diseases, inflammation, etc. | Hospital, Clinic, Laboratory | Whole Blood, Serum, Plasma, Urine | 10 µL | 3 min (HbA1c) | IFCC(mmol/mol)/NGSP/DCCT (%) values | collect blood samples | Add sample, Insert test kit, click 'Measure ' icon, Results printed automatically | Analyzer, User manual, adapter, data cable, QC card, printing paper, Accessory bag. Test kit: 25 individually packed testing cards, product specifications, one SD card and other consumables. |  | Quick quality control with QC card |  |  |  | 2.0 –14.0% | NA | NA | NA | USB, LAN |  |  |  | 7 inch touch screen |  | https://www.getein.com/uploadfile/attachment/48c00e3934dd037f05e7f92c4369b778.pdf | 1000 | NA | NA | NA | Available | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | Raj Biosis https://www.rajbiosis.in/ | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | INR 18000 | NA | NA | NA | NA | NA | NA | NA
28 | anuPath nanoA1c | Path Shodh Healthcare |  |  |  | ISO 13485 | [10 empty] | Immunoassay |  |  |  | Test strips |  |  |  | Handheld | 90 x 60 x 16 | 0.0716 | Blood glucose | anuPath is an easy to use point-of-care handheld diagnostic device based on cutting edge patented technology. It provides reliable and accurate diagnostic test results. | Professional use | Capillary whole blood | 10 µL | 1 min | IFCC(mmol/mol)/NGSP/DCCT (%) values | Collect blood sampl | Insert strip, select test, Apply sample, Get result | anuPath meter, Test strips, Lancing pen, Lancet, Alcohol swab |  |  |  | 10 - 40°C |  | 5.0 –15.0% | NA | NA | NA | USB | 15–40°C |  | One(1) CR 2032 3.0 V coin cell | 5 inch touch screen |  | https://pathshodh.com/pathshodh-anuPath-glucose-HbA1c-meter | NA | NA | NA | NA | Available | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | Pathshodh Healthcare Private Limited https://pathshodh.com/pathshodh-anuPath-glucose-HbA1c-meter | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | INR 9999 | NA | NA | NA | NA | NA | NA | NA
29 | Standard F100 | SD Biosensor, Korea | CE-IVD | [8 empty] | IFCC | NGSP |  |  |  | Immunoassay |  |  | Cartridge |  |  |  | Benchtop |  | 105x 135x 100 | 0.7 | Respiratory disease, Vector-borne disease, Hepatitis, Cardiac marker, Inflammation, Chronic disease, Hormone, Tumor marker. | STANDARD F100 Analyzer is a small fluorescence immunoassay device that can perform qualitative and quantitative analyses on infections, respiratory diseases and chronic diseases, and it is operable with battery. | Professional use | Capillary whole blood, anticoagulated venous whole blood | 5 µL | 3 min | IFCC(mmol/mol)/NGSP/DCCT (%) values | NA | NA | Analyzer, Calibration set, User manual, Quick Manual, External printer, Power cable, AC/DC adaptor. Test kit: Test device, buffer bottle, centrifuge tube |  | 2 Level controls (SDB HbA1c Control ) 10T/Kit |  | (-20 °C to 50 °C for analyzer) 2-30 °C for test kit |  |  |  |  | NA | Bluetooth |  |  | AC/DC adapter or battery | Graphic LCD |  | https://www.sdbiosensor.com/product/product_view?product_no=178 | NA | NA | NA | Available | Available | NA | NA | NA | NA | NA | NA | NA | NA | NA | https://www.komachine.com/en/companies/sd-biosensor/models | SD Biosensor Healthcare Pvt.ltd, Plot No. 63, Udyog Vihar, Phase-IV, Gurugram- 122015 (Haryana), India https://www.sdbiosensor.co.in/ | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
30 | Standard F200 | SD Biosensor, Korea | CE-IVD | [8 empty] | IFCC | NGSP |  |  |  | Immunoassay (Fluroscent) |  |  | Cartridge |  |  |  | Benchtop |  | 200 x 240 x 205 | 2.5 | Respiratory disease, Vector-borne disease, Hepatitis, Cardiac marker, Inflammation, Chronic disease, Hormone, Tumor marker. | STANDARD F200 Analyzer is a fluorescence immunoassay device that can perform qualitative and quantitative analyses on infections, respiratory diseases and chronic diseases, and it supports convenient user interface and LIS/HIS connectivity | Professional use | Capillary whole blood, anticoagulated venous whole blood | 5 µL | 3 min | IFCC(mmol/mol)/NGSP/DCCT (%) values | Collect blood, add on the test device, add 3 drops of buffer. | Insert test device and analyzer | Analyzer, PU protection film, Calibration set, User manual, Power cable, AC/DC adaptor. Test kit: Test device, buffer bottle, centrifuge tube | Lancet, capillary tube | 2 Level controls (SDB HbA1c Control ) 10T/Kit |  | (-20 °C to 50 °C for analyzer) 2-30 °C for test kit |  |  |  | < 3% | NA | LIS/ HIS |  |  | AC/DC adapter or battery | 7” Color touch screen |  | https://www.sdbiosensor.com/product/product_view?product_no=179 | NA | NA | NA | Available | Available | NA | NA | NA | NA | NA | NA | NA | NA | NA | https://www.komachine.com/en/companies/sd-biosensor/models | SD Biosensor Healthcare Pvt.ltd, Plot No. 63, Udyog Vihar, Phase-IV, Gurugram- 122015 (Haryana), India https://www.sdbiosensor.co.in/ | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
31 | APOTI™ Fluorescence Immunoassay Analyzer | APOTI | CE-IVD | FDA | [12 empty] | Immunoassay (Fluroscent) | [5 empty] | Cassette | Benchtop |  | 240 x 130 x 280 | <2.5 | Inflammation, Tumor, Nephrology, Cardiovascular, Coagulation, Endocrinology, Autoimmunity, Infectious disease | APOTI TM Fluorescence Immunoassay Analyzer is an analyzer based on the detection of fluorescence emitted during an immunoassay with antigen-antibody interaction. Fluorescence Immunoassay Analyzer offers the advantages of high accuracy, good stability, and fast results. | Professional use | Whole blood | 10 µl | 5 min | IFCC(mmol/mol)/NGSP/DCCT (%) values | Collect drop of blood on the test strip | Apply the test strip to the Apoti POCT analyzer after incubation and press test button | Analyzer, User manual, Adapter, power cable Test kit: 25 devices/tests |  |  | ID cards | 2-30 ° C for test kit, -10° C - 40° C for analyzer |  | 4.0 –14.5% | NA | NA | NA | RS232, USB (4 ports) and LAN interfaces | 4°C -30 °C |  | DC 12V | 7˝ LCD touch screen | 8000 | https://ivydiagnostics.com/wp-content/uploads/2021/07/Apoti-Sell-Sheet-b-002.pdf | NA | Available | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | Medistore https://www.medistorebd.com/ | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | Analyzer: BDT 55,000 Test kit: BDT 4,500. | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
32 | Fluorecare | Microprofit Biotech | CE-IVD |  |  | ISO 13495 | [6 empty] | NGSP |  |  |  | Immunoassay (Fluroscent) | [5 empty] | Cassette | Benchtop |  | 200×346×144 mm (WDH) | 2.5 | Thyroid, Fertility, Covid-19, Tumor marker,Gestrointestinal, Gestational, Bone metablism, Anemia, Inflammation, Cardiovascular, Autoimmune | Fluorecare is a quantitative analyzer, an advanced on-site medical device intended to deliver accurate testing results in minutes from 2-3 drops of human specimens. | Professional use | Whole blood | 10- 70 µl | 12 min | IFCC(mmol/mol)/NGSP/DCCT (%) values | The sample should be equilibrated to room temperature. Mix the sample well before running the test. Add the sample to the diluent and mix well | Apply sample to the test device, Insert into the analyzer and click "test" | Test kit: Instruction manual, ID card, diluent, foil pouch, test card |  |  | Automatic calibration through chip | At 15-25 degree |  | 4.0 –16% | NA | <10% | NA | 4 USB ports, 1 LAN port, 1 LIS port and 1 PC connection port |  |  | 100-240VAC, 50-60Hz |  |  | http://www.microprofit-bio.com/products/Analyzer.html | NA | Available | NA | NA | Available | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 1. "Invitro Diagnostics system", Mohanlalganj, Lucknow, Uttar Pradesh. 08045909313 2. "Diagnosol Healthcare", Jhotwara, Jaipur, Rajasthan 3. "TCM Healthcare Pvt.Ltd.28/2917", Aiswarya, Ernakulam, Kadavanthra, Kochi – 682020, Phone : +91 6282017736 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 1.INR 1,45,000 https://www.indiamart.com/proddetail/fluorecare-2851155604655.html?srsltid=AfmBOoqB-DkVvfOmk23vlP43IeZv2tfsZpGQ0H35H7dYvb4lDhjtAavY 2. INR 1,20,000 https://www.indiamart.com/proddetail/fluorecare-poct-immunoassay-analyzer-2854181472762.html?srsltid=AfmBOooRL4xYgjByd1URaPWH_Q4klbUIR1iYkGr1mJ-7lEoYDDf-BD7Z | NA | NA | NA | NA | NA | NA | NA
33 | AIFA-100 | Allcare | CE-IVD | [7 empty] | CDSCO | [5 empty] | Immunoassay (Time resolve immuno-fluorescence) | [5 empty] | Cassette | Benchtop |  | 260 x230x 140 | 2 | Bone metabolism, Cardiac, Thyroid, Inflammation, Fertility hormones, kidney marker | ALLCARE AIFA 100 is a POCT (Point of Care Test) System based on Time Resolve Immuno-Fluorescence Technology. It is an ideal system for quick and reliable diagnosis of biomarkers such as Cardiac, Cancer, Hormone (Thyroid & Fertility), Diabetes, Bone Metabolism and Infection diseases | Professional use | Serum, Plasma, Whole blood | 4 - 150 µl | 20min | NGSP (%) | NA | NA | Analyzer, Analyzer power cable, Analyzer power adapter, Incubator, Incubator power cable, Incubator adaptor, Paper roll, Analyzer manual. Test: Buffer, IC card for calibration, cassette |  |  | IC card | 2-30 ° C, 18 months |  | 2.0- 16% | NA | NA | NA | USB & ethernet port for LIS/HIS |  |  |  | 7˝ LCD touch screen | 100000 | https://www.alliedbiotech.co.in/allcare-aifa-100-ifa-analyzer.php | NA | NA | NA | NA | Available | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | "S.M. Infrastructure private limited", Tukoganj, Indore, Madhya Pradesh. Ph.no: 08046045615 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | INR 95000, https://www.indiamart.com/proddetail/aifa-100-25534321662.html?srsltid=AfmBOoqrLAMQC1ZMDo-tfLfhok83wqy0rlzxV_2lhXPyj_dBTGqlKcCJ | NA | NA | NA | NA | NA | NA | NA
34 | Unicorn Prima | Vector | CE-IVD |  |  |  |  | FSC |  |  |  |  | NGSP |  | CFDA |  | Lateral flow chromatography (Immunofluorescence) |  |  | Cartridge |  |  |  | Benchtop |  | 261 x 241x 115 (DxWxH) | 2 | Cardiac markers, Inflammation, Renal function, Hormone, Bone | The Unicorn Prima HbA1c Analyzer provides highly efficient and accurate testing using advanced fluorescence immunoassay technology. Its rapid, one-step testing process and user-friendly Android interface make it ideal for clinics and laboratories. | Professional use | Whole blood, plasma, serum, urine | 10 µl | 5 min | IFCC(mmol/mol)/NGSP/DCCT (%) values | NA | Insert test card and click start icon |  |  |  | NA | NA | NA | 2.0-14% | NA | NA | NA |  | 10-35°C | ≤ 70% | 100-240VAC, 50-60Hz | 7 inch touch screen | 10000 data | https://www.vectorbiotekindia.com/products_details.php?c_id=19 | NA | NA | NA | NA | Available | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 1."Manglam Traders", Triveni Nagar, Lucknow, Uttar Pradesh. Ph.no: 08047627996 2. "Vector Biotek pvt.ltd" , 425, New GIDC, Kabilpore, Navsari - 396424, Gujarat, India. Ph.no: 18002701027. http://www.vectorbiotekindia.com/ | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | INR 95000 https://www.indiamart.com/proddetail/unicorn-prima-immunofluorescence-quantitative-analyzer-23110860848.html?srsltid=AfmBOoqN8Yuy8m01cnR98OlTLX2IyFOjrktnSg94uxyLnVxX2o33s-Bq | NA | NA | NA | NA | NA | NA | NA
35 | SimplexTAS 101 | Tascom | CE-IVD | [8 empty] | IFCC |  |  |  |  | Immunoassay |  |  | Cartridge |  |  |  | Benchtop |  | 265 x 217 x 262 (WxLxH) | 3.4 | Renal function, Lipid profile, Liver function, Fatigue, Infection, Hematology, Gout | SimplexTAS is a new generation of Immuno-chemistry systems that quantitatively measure multiple parameters. It is a compact and portable point-of-care system that can use a drop of blood from the fingertip. | Professional use | Whole blood | 4 µl | 10 min | IFCC(mmol/mol)/NGSP/DCCT (%) values | Equilibrate the reagent for at least one hour before the test. Place the reagent cartridge in heating blocks for 20 minutes. Mix the cartridge well before use, collect the blood sample using capillary tube and place it in the cartridge | Turn on the analyzer one hour before the tests to warm it up. Place the cartridge in the rotor balance and close the lid. | Analyzer, Adapter, power cable, user manual. Test: Cartridges reagent |  |  |  | Analyzer 4-40° C (non-condensing) Test catridge 2-8 ° C for 12 months |  | 4.0-15.0%/ 20-140 mmol/ mol |  | NA | NA | USB 2.0 | 18–32°C | 10-70% (non- | 100-240 VAC 50-60Hz 0.75A | TFT LCD 5 inch touch panel | 2000 patient results/ 200 control results/ 100 IQC results | http://www.tascom.org/ | NA | NA | NA | Available | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | http://www.tascom.org/ | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA

## Abbreviations
Abbreviations
CE-IVD | Invitiro Diagnostics Certifications for Europe
CLIA | Clinical Laboratory Improvement Amendments
DCCT | Diabetes Control and Complications Trial
FDA | Food and Drug Administration
IFCC | International Federation of Clinical Chemistry and Laboratory Medicine
NGSP | National Glycohemoglobin Standardization Program
ACR | Albumin to creatinine ratio
CRP | C- reactive protein
JDS | Japan Diabetes Society
NMPA | National Medical Products Administration
CDSCO | Central Drugs Standard Control Organization
FSC | Food Safety Commission
KGMP | Korean Good Manufacturing Practice
HSA | Health Science Authority (Singapore)
PTS Diagnostics | Polymer Technology Systems